## Slide 1: Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies

## Slide 2: Module 3 Outline

## Slide 3: Module 3 Learning Objectives
At the end of Module 3, learners will be able to:
  Define key terms, such as calibration drift (CD), relative accuracy (RA), span value, etc.
  Compare performance specification (PS) and quality assurance (QA) procedures by pollutant, where relevant
  List relevant QA procedures by PS
  Provide examples of technologies that can be used for each PS
  Recognize why predictive emission monitoring systems (PEMS) PS are different from others

## Slide 4: Introduction to Performance Specifications
Performance specifications are used for evaluating the acceptability of continuous monitoring systems (CMS) at the time of, or soon after, installation and whenever specified in the regulations
Performance specifications are published in the Code of Federal Regulations (CFR), under Title 40 CFR Appendix B to Part 60
There are a total of 18 performance specifications. Most, but not all are pollutant specific
For those performance specifications that are pollutant-specific, this module covers commonly used technologies for analysis as well as QA

## Slide 5: Introduction to Performance Specifications (Cont’d)
Performance specifications are not designed to evaluate the installed CMS performance over an extended period of time.
Ongoing QA requirements are covered under 40 CFR Part 60, Appendix F.
The source owner or operator is responsible for calibrating, maintaining, and operating the CMS properly.
It should be noted that in many cases, the definitions, installation, and measurement location specifications, calculations and data analysis, and reference for the PS are the same as in PS-2, which we provide details for later in this presentation.
If these details differ, specific details will be provided for the relevant PS.

## Slide 6: Introduction to Quality Control and Quality Assurance
1. Quality Control
Quality control (QC) is the procedures, policies, and corrective actions necessary to ensure product quality.
2. Quality Assurance
Quality assurance (QA) procedures are used to evaluate the effectiveness of QC and the quality of data produced by any CMS that are used for determining compliance with the emission standards on a continuous basis as specified in the applicable regulation.
These procedures are pollutant-specific and published in Appendix F of 40 CFR 60.

## Slide 7: Technology Neutral
The majority of the performance specification and QA procedures do not specify the use of a specific measurement technology or are technology neutral.
This means that any sampling system using any technology that can, after being installed at the sampling location, pass the requirements of the PS and QA procedures is acceptable.

## Slide 8: To Be Covered….
PS | Pollutants Covered | QA Procedure
PS-1 | Opacity – Continuous Opacity Monitoring Systems (COMS) | Procedure 3
PS-2 | Sulfur Dioxide (SO2 ) and Oxides of Nitrogen (NOX) | Procedure 1
PS-3 | Oxygen (O2) and Carbon Dioxide (CO2) | Procedure 1
PS-4, 4A and 4B | Carbon Monoxide (CO) for PS-4 and 4A; and CO and O2 for PS-4B | Procedure 1
PS-5 | Total Reduced Sulfur (TRS) | 
PS-6 | Flow Rate – Continuous Emission Rate Monitoring Systems (CERMS) | 
PS-7 | Hydrogen Sulfide (H2S) | 
PS-8 | Volatile Organic Carbon (VOC) | 
PS-8A | Total Hydrocarbons (THC) | 
PS-9 | Gas Chromatography (GC) | 
PS-11 | Particulate Matter (PM) | Procedure 2
PS-12A and 12B | Mercury (Hg) | Procedure 5
PS-15 | Fourier Transform Infrared (FTIR) | 
PS-16 | PEMS | 
PS-18 | Hydrogen Chloride (HCl) | Procedure 6

## Slide 9: Setting the Stage: Commonly Used Terms

## Slide 10: Setting the Stage: Calibration Error and Calibration Drift

## Slide 11: Setting the Stage: Response Time and Out-of-Control

## Slide 12: Quality Assurance Procedure Audits
Audit procedures are critical for verifying proper performance of the monitoring systems and identifying problems which may lead to inaccurate emissions accounting.
There are four main types of audits discussed in QA procedures:

## Slide 13: What is a Relative Accuracy Test Audit?
The ANNUAL comparative evaluation of the CEMS performance using a RM
Consists of:
  9 or more RM test runs
  Usually 21 minutes in duration

## Slide 14: What is a Cylinder Gas Audit?
Usually performed QUARTERLY – in three of four quarters, annually
  With RATA conducted in the fourth quarter
Gases needed and methodology used are found in applicable QA procedure
Audit gases must be certified by or traceable to National Institute of Standards and Technology (NIST)

## Slide 15: What is a Calibration Drift Assessment?
The DAILY check of the difference in the CEMS readings from a known value, usually a calibration gas
Performed to demonstrate the stability of the CEMS calibration – how does it fluctuate over time?
Initial certification usually requires a 7-day drift test
Daily drift test required for ongoing operation

## Slide 16: What is a Relative Accuracy Audit?
An alternative QUARTERLY audit procedure which correlates the CEMS data to simultaneously collected RM data
Performed like a RATA, but only requires three RM test runs
Not used very often, but is an option

## Slide 17: Performance Specification 1 (PS-1) Specifications and Test Procedures for Continuous Opacity Monitoring Systems (COMS) in Stationary Sources
17

## Slide 18: PS-1 Requirements
PS-1 contains requirements for:
  Manufacturers
  Owners/operators
    Installation
    Performance test requirements
Photo reprinted with permission from Alabama Department of Environmental Management

## Slide 19: Requirements for Manufacturers
Opacity manufacturers must comply with a comprehensive series of design and performance specifications and test procedures to certify opacity monitoring equipment before shipment to the end user.

## Slide 20: Requirements for Owners/Operators: Installation
Install COMS at a location where the opacity measurements are representative of the total emissions from the affected source:
  4 duct diameters downstream from any disturbance
  2 duct diameters upstream from any disturbance
  Condensed H2O vapor is not present
  NOTE: Additionally, installed COMS must be accessible for maintenance.

## Slide 21: Requirements for Owners/Operators: Performance Tests
Calibration Error Check
  A three-point CE check
  Criteria: <3% opacity
Optical Alignment Assessment
  Verify and record that all alignment indicator devices show proper alignment
  Criteria: Is it aligned?

## Slide 22: Requirements for Owners/Operators: Performance Tests (Cont’d)
System Response Time
Measure the amount of time needed for a 95% step change in the COMS data recorder
Criteria: <10 sec
Averaging Period Calculation and Recording Check
Following the CE check, conduct a check of the averaging period calculation and recording
Criteria:
  Averaging period check + 2% opacity
  Data recorder: resolution < 0.5% opacity

## Slide 23: Operational Test Period
Total time of 168 hours (7-day drift test) at normal operation
  Includes shut-downs, if normal occurrences, but total operating time during the test must be 168 hours
  Batch cycles must include at least one full batch cycle

## Slide 24: Procedure 3 – Quality Assurance Requirements for Continuous Opacity Monitoring Systems at Stationary Sources
What are the basic functions of Procedure 3?
Assessment of the quality of your COMS data
Control and improvement of the quality of your COMS data by implementing QC requirements and corrective actions
Requires:
  Daily instrument drift checks
  Status indicator checks
  Quarterly performance audits
    Optical alignment
    CE
    Zero compensation
  Annual zero alignment

## Slide 25: Procedure 3 - Limitations
Opacity cannot be measured accurately in the presence of condensed water vapor. Thus, COMS opacity compliance determinations cannot be made when condensed water vapor is present, such as downstream of a wet scrubber without a reheater or at other saturated flue gas locations.

## Slide 26: Procedure 3 – Required Quality Control Program
A QC program must, at a minimum, include written procedures which describe in detail complete step-by-step procedures and operations for these activities:
  Performing drift checks,
  Performing quarterly performance audits,
  Checking the zero alignment of the COMS, and
  Corrective action for a malfunctioning COMS.
It is required to keep the QA/QC written procedures on site and available for inspection by Federal, state, and/or local enforcement agencies.

## Slide 27: Procedure 3 – Auditing Requirements

## Slide 28: Procedure 3 – Limits for Excessive Inaccuracy
Your COMS is out-of-control if either the zero drift check or upscale drift check exceeds twice the applicable drift specification in PS-1 for any one day.
Your COMS is out-of-control if the zero alignment error exceeds 2 percent opacity.
Your COMS is out-of-control if the results of a quarterly performance audit indicate noncompliance with the following criteria:
  The optical alignment indicator does not show proper alignment,
  The zero compensation exceeds 4 percent opacity, or
  The calibration error exceeds 3 percent opacity.
Note: You must adhere to the data capture criterion specified in the applicable subpart.
Excessive zero or upscale drift
Excessive zero alignment
Quarterly performance audit

## Slide 29: Procedure 3 – Corrective Actions
A corrective action program must be in place to address the repair and/or maintenance of your COMS. The corrective action program must:
  Address routine/preventative maintenance and various types of analyzer repairs, and
  Establish what diagnostic testing must be performed to ensure that the COMS is collecting valid, quality-assured data.
  NOTE: Recommended maintenance and repair procedures and diagnostic testing after repairs may be found in an associated guidance document.
  https://www3.epa.gov/ttn/emc/perfspec/suggested_COMS_diagnostic_tests.pdf

## Slide 30: Opacity Transceiver

## Slide 31: Commonly Used Technology - Transmissometry

## Slide 32: Double Pass Transmissometer
Courtesy of Janke 1979

## Slide 33: Double Pass Transmissometer (Cont’d)
Courtesy of Janke 1984

## Slide 34: 1. Let’s Test Your Knowledge!
34

## Slide 35: PS-1 and Procedure 3
What are the requirements of the CD test under PS-1?
Check Answer
1. Procedures for performing drift checks must include both zero and upscale drift and the status indicators check.
2. How is system response time determined?
Check Answer
2. System response time is determined by measuring the amount of time needed for a 95% step change in the COMS data recorder.
3. What is the OOC period for each audit in procedure 3?
Check Answer
3. Daily Audit - if the zero drift check or upscale drift check exceeds twice the applicable drift specification in PS-1 for any one day
Quarterly audit - if:
  The optical alignment indicator does not show proper alignment,
  The zero compensation exceeds 4% opacity, or
  The CE exceeds 3% opacity
Annual audit - if the zero-alignment error exceeds 2% opacity.
4. What are the responsibilities of the opacity monitor manufacturer?
Check Answer
4. Opacity manufacturers must comply with a comprehensive series of design and performance specifications and test procedures to certify opacity monitoring equipment before shipment to the end user.

## Slide 36: Performance Specification 2 (PS-2) Specifications and Test Procedures for Sulfur Dioxide (SO2) and Oxides of Nitrogen (NOX) in Stationary Sources
36

## Slide 37: Overview of PS-2
Evaluates the acceptability of SO2 and NOX CEMS at the time of installation or soon after
    The CEMS may include, a diluent (O2 or CO2) monitor
Includes installation and measurement specifications as well as requirements for:
    7-day CD test
    RATA
NOTE: PS-2 serves as the framework for most other performance specifications.
Image courtesy of Thermo Fisher Scientific™

## Slide 38: CEMS Installation and Measurement Location Specifications
Must be accessible and representative.
At least two equivalent diameters downstream and one-half an equivalent diameters upstream from any flow disturbance
Not required that RM sampling location to be the same as CEMS location

## Slide 39: Calibration Drift Test

## Slide 40: Relative Accuracy Test Audit

## Slide 41: Determining if Stratification Exists
  Calculate the mean value of all the sample points
  Find the difference between the mean value and each individual sample value
      If the mean pollutant concentration is more than 10% different from any single sample point, then stratification exists
            Must use the points located at 16.7, 50.0, and 83.3 % of the entire measurement line
  Conduct all necessary RM tests within 3 cm (1.2 in.) of the traverse points, but no closer than 3 cm (1.2 in.) to the stack or duct wall

## Slide 42: Reference Method Traverse Points Minimum Requirements
Establish a measurement line through the stack centroid that includes three traverse points at:
16.7,
50.0, and
83.3% of the duct/stack diameter
If a measurement line is longer than 2.4 m (7.8 ft) and stratification is not expected, then use traverse points at:
0.4,
1.2, and
2m from stack/duct wall for the traverse points.
Note: This option cannot be used with a wet scrubber or at points where two streams with different pollutant concentrations combine.

## Slide 43: Relative Accuracy Test Audit Procedure
Number of RM Tests:
  Conduct a minimum of nine test runs
  Data from all test runs must be reported, including the rejected runs
Calculate:
  Mean difference between the RM and CEMS values in the units of the emission standard,
  Standard deviation,
  Confidence coefficient, and
  RA according to the procedures in section 12.0.

## Slide 44: Relative Accuracy Performance Criteria
 | Calculate… | Relative Accuracy Criteria
If average emissions during the RATA are >50% of emission standard | Use Eq. 2-6, with RM in the denominator | <20%
If average emissions during the RATA are <50% of emission standard | Use Eq. 2-6, emission standard in the denominator | <10%
For SO2 emission standards < 130 but > 86 ng/J (0.30 and 0.20 lb/million Btu) | Use Eq. 2-6, emission standard in the denominator | <15%
For SO2 emission standards <86 ng/J (0.20 lb/million Btu) | Use Eq. 2-6, emission standard in the denominator | <20%

## Slide 45: Procedure 1 – Quality Assurance Requirements for Gas CEMS Used for Compliance Determination
What are the basic functions of Procedure 1?
Evaluates the effectiveness of QA/QC procedures and the quality of data produced by any CEMS used for determining compliance
Specifies the minimum QA requirements necessary for the control and assessment of the quality of CEMS data submitted
Consists of two distinct and equally important functions:
  The assessment of the quality of the CEMS data by estimating accuracy
  The control and improvement of the quality of the CEMS data by implementing QC policies and corrective actions

## Slide 46: Procedure 1 – Checks/Auditing Requirements
Daily Checks
Quarterly Checks
CD at two concentration values – zero and high-level
CGA – if applicable, may be conducted in three of four calendar quarters, but in no more than three quarters in succession
RAA – may be conducted three of four calendar quarters, but in no more than three quarters in succession
Other Alternative Audits – may be conducted as approved by the administrator for three of four calendar quarters
Annual Checks
RATA – conducted at least once every four calendar quarters except for other alternative audits

## Slide 47: Procedure 1 – Performance Criteria
Calibration Drift
Cylinder Gas Audit
Must not exceed twice the applicable drift specification found in Appendix B for five consecutive days, or four times the applicable drift specification in Appendix B (i.e., 2.5%) on any one day. If so, the CEMS is considered OOC
Must be less than + 15% of the average audit value or + 5 ppm, whichever is greater, or the CEMS is OOC
Relative Accuracy Test Audit
Same as the RA requirement in the applicable PS or the CEMS is OOC

## Slide 48: Commonly Used Technologies: Pulsed Fluorescence and Chemiluminescence Analyzer
Uses the property of SO2 molecules to absorb ultraviolet (UV) light and become excited at one wavelength, then decay to a lower energy state emitting UV light at a different wavelength, the measured emitted light corresponding to the concentration of SO2 in the sample gas
    The pulsing of the UV source lamp allows the analyzer to use both the light and dark phases of the pulsed light to continuously detect and correct for electronic noise, and to measure lower pollutant concentrations.
Uses the light-emitting chemical reaction of NO and analyzer-generated ozone to measure the concentration of the NO in a gas sample. A successive measurement of the NO, plus NO converted from the NO2 in the sample, gives a total NOx measurement; the difference between the two measurements is equal to the NO2 concentration in the sample.
Pulsed Fluorescence
Chemiluminescence Analyzer

## Slide 49: SO2 Analyzer
Microprocessor control
SO2 specific
Reflective UV filtering
Hermetically sealed UV lamp
No consumables
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 50: SO2 Monitor
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 51: Chemiluminescence NOx Analyzer
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 52: How is NOX Measured?
Chemiluminescence Technique
Intensity of emitted light is proportional to NO concentration

## Slide 53: NOX Monitor
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 54: 2. Let’s Test Your Knowledge!
54

## Slide 55: PS-2 – Question 1
1. During the PS-2 seven-day CD test, most plants must be operating more than _______% of normal plant load.
A. 50
B. 75
C. The plant does not have to be operating for this test.
D. No operating load is specified.
1. D. No operating load is specified.
55

## Slide 56: PS-2 – Question 2
2. During preliminary testing, you see the mean pollutant concentration has a 25% difference from some of the sample points. Is the source stratified or not stratified?
A. Stratified
B. Not stratified
2. A. The source is stratified.
56

## Slide 57: PS-2 – Question 3
3. The relative accuracy test audit, or RATA, must be passed prior to installation of the CEMS.
True
False
3. False. The CEMS must be installed before certification testing can begin.

## Slide 58: Procedure 1
Name three types of quarterly audits.
Check Answer
1. CGA – if applicable, may be conducted in three of four calendar quarters, but in no more than three quarters in succession
RAA – may be conducted three of four calendar quarters, but in no more than three quarters in succession
Other Alternative Audits – may be conducted as approved by the Administrator for three of four calendar quarters
2. When should a RATA be performed?
Check Answer
2. One RATA is required at least every four calendar quarters, except where the affected facility is off-line (does not operate) in the fourth calendar quarter since the quarter of the previous RATA. If this occurs, the RATA shall be performed in the quarter in which the unit recommences operation.
3. What are the QC requirements of procedure 1?
Check Answer
3. At a minimum, each QC program must include written procedures in detail on step-by-step procedures for each of the following:
  Calibration of CEMS.
  CD determination and adjustment of CEMS.
  Preventive maintenance of CEMS (including spare parts inventory).
  Data recording, calculations, and reporting.
  Accuracy audit procedures including sampling and analysis methods.
  Program of corrective action for malfunctioning CEMS.
58

## Slide 59: Performance Specification 3 (PS-3) Specifications and Test Procedures for Oxygen (O2) and Carbon Dioxide (CO2) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
59

## Slide 60: Overview of PS-3

## Slide 61: Calibration Drift and Relative Accuracy Performance Criteria – PS-3

## Slide 62: Quality Assurance Procedure 1
PS-2, 3, and 4 use QA procedure 1.
For more information, review slides under PS-2.

## Slide 63: Commonly Used Technologies for PS-3
A zirconium oxide O2 analyzer is an electrochemical cell which is porous to O2 when heated to high temperature, allowing the O2 to pass from the high concentration side (reference) to low concentration side (sample) and generating a voltage proportional to the difference in O2 concentrations.

## Slide 64: Commonly Used Technologies for PS-3: O2 & CO2 Analyzer
O2 & CO2 Analyzer
Photo reprinted with permission from Alabama Department of Environmental Management

## Slide 65: Commonly Used Technologies for PS-3: Paramagnetic O2 Analyzer and Non-Dispersive Infrared

## Slide 66: Performance Specification 4 (PS-4) Specifications and Test Procedures for Carbon Monoxide (CO) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
66

## Slide 67: Overview of PS-4

## Slide 68: Calibration Drift and Performance Specification Criteria – PS-4

## Slide 69: Quality Assurance - Procedure 1
PS-2, 3, and 4 use QA procedure 1.
For more information, review slides under PS-2.

## Slide 70: Commonly Used Technology for PS-4
The gas filter correlation (GFC) method is similar to NDIR, but instead of two parallel cells it uses a rotating filter wheel with a section of the wheel containing a sample of the gas being measured.
IO
I
Mirrors
Modulator
Gas filter cell (CO)
Neutral filter (N2)
Sample
gas
Mirror
Infrared source
Filter
Detector

## Slide 71: 3. Let’s Test Your Knowledge!
71

## Slide 72: PS-3
1. For PS-3, CEMS calibration must not drift by more than ______ O2 or CO2 from the reference value.
A. 0.02%
B. 0.5%
C. 0.1%
D. 1%
1. B. For PS-3, CEMS calibration must not drift by more than 0.5% O2 or CO2 from the reference value.
2. Name three commonly used O2/CO2 analyzer technologies.
Check Answer
2. Three commonly used O2/CO2 analyzer technologies are paramagnetic, zirconium oxide, and NDIR.
3. The results of a RATA are considered valid if the calculated relative accuracy is less than ____ , or if the RA is less than or equal to _____ when the applicable emission standard is used.
A. 20.0% and 1.0%
B. 15.0% and 5.0%
C. 20.0 ppm and 5.0 ppm
D. 15.0% and 1.0%
3. A. The results of a RATA are considered valid if the calculated relative accuracy is less than 20.0%, or if the RA is less than or equal to 1.0% when the applicable emission standard is used.

## Slide 73: PS-4 – Question 1, 2
1. For PS-4, CEMS calibration must not drift by more than _____CO from the reference value.
A. 1%
B. 1.5%
C. 3.5%
D. 5%
1. D. For PS-4, CEMS calibration must not drift by more than 5% CO from the reference value.
2. For PS-4, the results of a RATA test are considered valid if the calculated relative accuracy is less than ____, or if the RA is less than or equal to ____ when the applicable emission standard is used.
A. 10% and 5%
B. 15% and 5%
C. 20% and 1%
D. 10% and 1%
2. A. For PS-4, the results of a RATA test are considered valid if the calculated relative accuracy is less than 10%, or if the RA is less than or equal to 5% when the applicable emission standard is used.

## Slide 74: PS-4 – Question 3, 4
3. Name two commonly used CO analyzer technologies.
Check Answer
3. Two commonly used CO analyzer technologies are IR and GFC.
4. PS-4 was primarily written to be used for CO analyzers with what span?
Check Answer
4. PS-4 was primarily written to be used for CO analyzers with a 1000 ppmv.

## Slide 75: Performance Specification 4A (PS-4A) Specifications and Test Procedures for Carbon Monoxide (CO) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
75

## Slide 76: PS-4A Overview
PS-4A is for evaluating the acceptability of CO CEMS at the time of installation or soon after and whenever specified in an applicable subpart of the regulations.
  The main difference between this PS and PS-4 is that it was developed primarily for CEMS that comply with low emission standards (less than 200 ppmv).

## Slide 77: PS-4A Requirements
Most aspects of this PS are the same as PS-2, with the exception of:
    CD performance criteria - same as PS-4, refer to previous slides
    RA performance criteria - same as PS-4, refer to previous slides
    Response time
        The CEMS response time shall not exceed 240 seconds to achieve 95% of the final stable value.
  Reference methods for PS-4A are 10, 10A, and 10B

## Slide 78: 4. Let’s Test Your Knowledge!
78

## Slide 79: PS-4A
1. For PS-4A, the results of a RATA test are considered valid if the calculated RA is less than ____, or if the RA is less than or equal to ____ when the applicable emission standard is used.
A. 10% and 5%
B. 10 ppm and 5 ppm
C. 20% and 1%
D. 10% and 1%
1. A. For PS-4A, the results of a RATA test are considered valid if the calculated RA is less than 10%, or if the RA is less than or equal to 5% when the applicable emission standard is used.
2. Name three reference methods that are associated with PS-4A.
Check Answer
2. Three reference methods associated with PS-4A are M10, 10A, and 10B.
3. PS-4A was primarily written to be used for CO analyzers with what span?
Check Answer
3. PS-4A was primarily written to be used for CO analyzers with less than 200 ppmv.

## Slide 80: Performance Specification 4B (PS-4B) Specifications and Test Procedures for Carbon Monoxide (CO) and Oxygen (O2) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
80

## Slide 81: PS-4B Overview
PS-4B is to be used for evaluating the acceptability of CO and O2 CEMS at the time of or soon after installation and whenever specified in the regulations.
The CEMS may include, for certain stationary sources:
  flow monitoring equipment to allow measurement of the dry volume of stack effluent sampled, and
  an automatic sampling system.

## Slide 82: Performance and Equipment Specifications: Data Recorder Scale: O2 and Data Recorder Scale: CO
The output range must include the full range of expected concentration values in the gas stream including the zero and span values
Span must be 25%; can be higher if O2 concentration at sampling point can exceed 25%
Must record all readings within a measurement range with a resolution of 0.5%
Data Recorder Scale: O2
Low – range span must be 200 ppm
High – range span must be 3000 ppm
Must record all readings within a measurement range with a resolution of 0.5%
Data Recorder Scale: CO

## Slide 83: Performance and Equipment Specifications: Calibration Drift and Calibration Error
O2: Same as PS-3
CO: Same as PS-4, except it must not drift by more than 3% of the span value
Calibration Drift
O2: Mean difference between CEMS and reference values at all 3 test points must be ≤ 0.5% O2
CO: Mean difference between CEMS and reference values at all 3 test points ≤ 5.0% of span
Calibration Error
NOTE: CE and response time tests should be conducted during the CD test period. Response time must not exceed 240 seconds.

## Slide 84: Performance and Equipment Specifications: Relative Accuracy: CO and O2
For O2, same as PS-3 and for CO, same as PS-4A
Relative Accuracy: CO and O2

## Slide 85: Alternative Relative Accuracy Procedures
Conduct complete CEMS status check per manufacturer’s written instructions
Instrument must pass CE and CD specifications and have administrator approval

## Slide 86: Quality Assurance Procedure 1 and Commonly Used Technology
PS-2, 3, and 4, 4A, and 4B use QA procedure 1. For more information, review slides under PS-2.
Commonly used technology is a GFC, same as PS-4

## Slide 87: 5. Let’s Test Your Knowledge!
87

## Slide 88: PS-4B – Question 1, 2
1. PS-4B is primarily used for testing what two pollutants together?
Check Answer
1. PS-4B is primarily used for testing O2 and CO together.
2. For PS-4B, the results of a RATA test are considered valid if the calculated RA is less than ____, or if the RA is less than or equal to ____ when the applicable emission standard is used.
A. 5.0% and 0.5%
B. 1.0% and 5.0%
C. 10.0% and 1.0%
D. 5.0% and 1.0%
2. A. For PS-4B, the results of a RATA test are considered valid if the calculated RA is less than 5.0%, or if the RA is less than or equal to 0.5% when the applicable emission standard is used.

## Slide 89: PS-4B – Question 3, 4
3. For PS-4B CE, the mean difference between CEMS and reference values at all 3 test points must be ≤ 0.5% for which pollutant(s)?
A. CO
B. O2
C. CO and O2
D. None of the above
3. B. O2.
4. What is one commonly used technology associated with PS-4B?
Check Answer
4. GFC is one commonly used technology associated with PS-4B.

## Slide 90: Performance Specification 5 (PS-5) Specifications and Test Procedures for Total Reduced Sulfur (TRS) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
90

## Slide 91: PS-5 Overview
Evaluating the applicability of TRS CEMS at the time of installation or soon after and whenever specified in an applicable subpart of the regulations.
Three reference methods, which are 16, 16A, and 16B.
The CEMS may include O2 monitors which are subject to PS-3.
  Most aspects of this PS are the same as PS-2, with the exception of:
    CD performance criteria
    RA performance criteria

## Slide 92: Calibration Drift and Relative Accuracy
Method Performance: Must perform CD test and relative RA test at time of initial installation or soon after.
Calibration Drift
Relative Accuracy
CEMS must not drift from the reference value of the calibration gas by no more than 5% of CEMS span value for 6 out of 7 days.
  No adjustments to CEMS prior to daily measurement
  Conduct CD test at 2 points (0-20% and 50-100% of full scale) (PS-2- 6.1.2)
  Must determine CD of diluent monitor separately (PS-3)
RA Test- RA of CEMS must be no greater than 20% of average RM or 10% when the applicable standard is used to calculate RA.
NOTE: There is no promulgated QA procedure for TRS monitors, however, subparts or regulations may require ongoing QA.

## Slide 93: Commonly Used Technology for PS-5
NOTE: Pulsed Fluorescence – see details under PS-2

## Slide 94: 6. Let’s Test Your Knowledge!
94

## Slide 95: PS-5
1. What are three reference methods that are commonly associated with PS-5?
Check Answer
1. The three reference methods that are commonly associated with PS-5 are 16, 16A, and 16B.
2. For PS-5, the results of a RATA test are considered valid if the calculated RA is less than ____ , or if the RA is less than or equal to ____ when the applicable emission standard is used.
A. 10% and 5%
B. 15% and 5%
C. 20% ad 10%
D. 10% and 1%
2. C. For PS5, the results of a RATA test are considered valid if the calculated RA is less than 20%, or if the RA is less than or equal to 10% when the applicable emission standard is used.

## Slide 96: Performance Specification 6 (PS-6) Specifications and Test Procedures for Continuous Emission Rate Monitoring Systems (CERMS) in Stationary Sources
96

## Slide 97: Overview of PS-6
Evaluating the acceptability of CERMS. Definitions are the same as those in PS-2 with the exceptions of:
  CERMS—the total equipment required for the determining and recording the pollutant mass emission rate (in terms of mass per unit of time); and
  Flow Rate Sensor—portion of the CERMS that senses the volumetric flow rate and generates an output proportional to that flow rate. The flow rate sensor shall have provisions to check the CD for each flow rate parameter that it measures individually (e.g., velocity, pressure).
The CD and RA tests are conducted to determine conformance of the CERMS to the specification.
Reference methods used for determining flow under PS-6 are 2, 2A, 2B, 2C, and 2D.
97

## Slide 98: Calibration Drift
Determined separately for each analyzer
  Shall not exceed 3% of the high-level value
Conduct the CD tests for pollutant concentration at the two values specified in section 6.1.2 of PS-2.
For other parameters (e.g., velocity, pressure, flow rate), use two analogous values (e.g., Low: 0-20% of full scale, High: 50-100% of full scale).

## Slide 99: Relative Accuracy
The RA of the CERMS shall be no greater than 20% of the mean value of the RM's test data in terms of the units of the emission standard, or 10% of the applicable standard, whichever is greater.

## Slide 100: Commonly Used Technologies for PS-6
Pitot tubes use the differential pressure between the measurements of total pressure and the static pressure at a point in the stack to calculate the stack gas velocity and volumetric flowrate.

## Slide 101: Commonly Used Technologies for PS-6 (Cont’d)
An ultrasonic flowmeter uses a pair of transmitter/receivers mounted on opposite sides of the stack, with one upstream from the other. The signal is alternated between them, sending it in the direction of stack gas flow, where it is speeded up, and then against the direction of flow, where it is slowed down. The difference in the time between the two signals is proportional to the stack gas velocity.
Acoustic Velocity

## Slide 102: 7. Let’s Test Your Knowledge!
102

## Slide 103: PS-6
1. What is PS-6 primarily used for to perform?
Check Answer
1. PS-6 is primarily used to perform CERMS installation and certification.
2. For PS-6, CERMS calibration must not drift by more than ___% from the reference value.
A. 3%
B. 5%
C. 2%
D. 0.5%
2. A. 3%
3. What reference methods are used for determining flow under PS-6?
Check Answer
3. Reference methods 2, 2A, 2B, 2C, and 2D are used for determining flow under PS-6.
4. When measuring flow rates, what are two commonly used techniques?
Check Answer
4. When measuring flow rates, pilot tubes and ultrasonic flow meters are two commonly used techniques.

## Slide 104: Performance Specification 7 (PS-7) Specifications and Test Procedures for Hydrogen Sulfide (H2S) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
104

## Slide 105: PS-7 Overview

## Slide 106: PS-7 Calibration Drift and Relative Accuracy
Calibration Drift
Relative Accuracy
CEMS must not drift from the reference value of the calibration gas by any more than 5% of CEMS span value for 6 out of 7 days
No adjustments to CEMS prior to daily measurement
Conduct CD test at 2 points (0-20% and 50-100% of full scale)
CEMS RA must be no greater than 20% of average RM or 10% of the applicable standard
RA Test - Perform minimum of 9 RM test runs, may do more, but only a maximum of 3 runs may be discarded. Must report 9 runs.
Use RM 11, 15, 16
  RM 11 sample run times shall be at least 10-minutes and (0.35 dscf or 0.010 dscm) and taken at 30-minute intervals
  RM 15 and 16 sample runs shall consist of 2 injections equally spaced over 30-minute period
  NOTE: Some aspects of this PS are the same as PS-2.

## Slide 107: Commonly Used Technology for PS-7
GC—see description in PS-5
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 108: 8. Let’s Test Your Knowledge!
108

## Slide 109: PS-7
1. What is PS-7 primarily used to certify?
Check Answer
1. PS-7 is primarily used for certifying H2S CEMS.
2. What reference methods are associated with PS-7?
Check Answer
2. Reference methods 11, 15, and 16 are associated with PS-7.
3. What is a commonly used technology that is associated with PS-7?
Check Answer
3. A GC is a commonly used technology associated with PS-7.

## Slide 110: Performance Specification 8 (PS-8) Specifications and Test Procedures for Volatile Organic Compounds (VOC) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
110

## Slide 111: PS-8 Overview
Must select the same measurement technology as the reference method or if not specified, the technology is based on knowledge of the source emissions.

## Slide 112: PS-8 Calibration Drift and Relative Accuracy
Calibration Drift
Relative Accuracy
< 2.5% of span value
No adjustments to CEMS prior to daily measurement
< 20% of average RM or < 10% of the applicable standard
Use RM as specified in applicable subpart or regulation
NOTE: Subparts or regulations may require on-going QA.

## Slide 113: Commonly Used Technology for PS-8
A flame ionization detector (FID) measures the current induced by ions attracted to and hitting a collector plate. The ions are formed by the combustion of organic compounds in a sample gas.
A gas sample is extracted from the source through a heated sample line and heated filter to an FID.
An FID measures the current, which is directly proportional to the concentration of VOC in the sample.
  Results are reported as volume concentration equivalents of propane.
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 114: 9. Let’s Test Your Knowledge!
114

## Slide 115: PS-8 – Question 1, 2
1. What is PS-8 primarily used to evaluate?
Check Answer
1. PS-8 is primarily used to evaluate CEMS that measures a mixture of VOC.
2. For PS-8, CEMS calibration must not drift by more than ___% from the reference value.
A. 5%
B. 2.5%
C. 8%
D. 10%
2. B. 2.5%

## Slide 116: PS-8 – Question 3, 4
3. For PS-8, the results of a RATA test are considered valid if the CEMS is no greater than ___ % of the mean value of the RM data, or if the applicable emission standard calculation is less than or equal to ____ %.
A. 10% and 5%
B. 15% and 5%
C. 20% and 10%
D. 10% and 1%
3. C. For PS-8, the results of a RATA test are considered valid if the CEMS is no greater than 20% of the mean value of the RM data, or if the applicable emission standard calculation is less than or equal to 10%.
4. What is the most commonly used technology associated with PS-8?
Check Answer
4. The most commonly used technology associated with PS-8 is an FID.

## Slide 117: Performance Specification 8A (PS-8A) Specifications and Test Procedures for Total Hydrocarbons (THC) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
117

## Slide 118: PS-8A Overview

## Slide 119: PS-8A Requirements
RATA and absolute calibration audits (ACA) are not required. The CD, CE, and response time tests are performed in lieu of a RA or ACA. If not passed they must be performed again until all are passed.
Calibration Drift Test Period
Calibration Error Test and
Response Time Test
While a unit is operating, determine the CD for seven consecutive operating days, make no adjustment to system prior to performing CD test.
The CEMS must not drift by +3ppm or +3% of span value after each 24-hour period of the 7-day drift test for both the zero and span gases.
Conduct the CE and response time tests during the CD test period.
The mean difference between the CEMS and reference values at all three test points must be no greater than 5 ppm (±5% of the span value).

## Slide 120: Commonly Used Technology for PS-8A
FID - see details under PS-8
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 121: 10. Let’s Test Your Knowledge!
121

## Slide 122: PS-8A
1. What is PS-8A primarily used to evaluate?
Check Answer
1. PS-8A is primarily used to evaluate one THC CEMS.
2. For a PS-8A calibration drift test, the CEMS must not drift by ____ppm or ___% of span value after each 24-hour period of the 7-day drift test for both the zero and span gases.
A. + 1 or + 1
B. + 4 or +4
C. + 5 or + 5
D. +3 or +3
2. D. For PS-8A calibration drift test, the CEMS must not drift by +3 ppm or +3 % of span value after each 24-hour period of the 7-day drift test for both the zero and span gases.

## Slide 123: Performance Specification 9 (PS-9) Specifications and Test Procedures for Gas Chromatography (GC) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
123

## Slide 124: PS-9 Overview
Applies to CEMS that use GC to measure a specific target list of organic compounds.
GC is defined as that portion of the system that separates and detects organic analytes and generates an output proportional to the gas concentration. The GC must be temperature controlled.
Definitions unique to this PS are included in the subsequent slides.
Calibration precision, CE, and performance audit tests are conducted to determine conformance of the CEMS with these specifications. Daily calibration and maintenance requirements are also specified.

## Slide 125: Sample Collection, Preservation, Storage and Transport
Installation and Measurement Location Specifications
  Install CEMS where measurements are representative of the source.
  Sampling location should be at least two equivalent duct diameters from control device, point of pollutant generation, or any point that might cause a change in pollutant concentrations.
Pre-Test Preparation Period
  Use the procedures in method 18 of 40 CFR Part 60, Appendix A to perform initial tests to determine the proper GC conditions that provide good resolution and minimum analysis time for the analytes of interest.

## Slide 126: Performance Audits and Calibration

## Slide 127: Calibration Error

## Slide 128: Commonly Used Technology for PS-9
GC—see description in PS-7
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 129: 11. Let’s Test Your Knowledge!
129

## Slide 130: PS-9
1. PS-9 applies to CEMS that use___.
Check Answer
PS-9 applies to CEMS that use GC to measure gaseous organic compound emissions.

## Slide 131: Performance Specification 11 (PS-11) Specifications and Test Procedures for Particulate Matter (PM) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
131

## Slide 132: PS-11 Overview
Establishes the initial installation and performance procedures that are required for evaluating the acceptability of a PM CEMS.
PS-11 requires initial installation and calibration procedures that confirm the acceptability of the CEMS when it is installed and placed into operation.
A site-specific correlation must be developed of the PM CEMS response against manual gravimetric RM measurements (including those made using EPA methods 5, 5I, or 17).

## Slide 133: Installation and Performance
Initial installation and performance may include:
  Diluent monitor - O2, CO2, or other monitors specified in applicable regulation.
  Auxiliary monitoring equipment for temperature, pressure, moisture content and/or volume of stack effluent.
  Automatic sampling system that measures in units of mass concentration.

## Slide 134: Drift and Correlation Test
Performance Criteria based on the following:
  7-day drift test
  Initial correlation test
  Sampling periods
  Cycle/Response time
7-Day Drift Test
  <2% of the upscale value (includes O2, CO2 monitors).
Initial Correlation Test
  Based on a technique of correlating PM CEMS responses relative to emission concentrations determined by the RM (EPA Method 5, 5I, 17). Unlike gaseous CEMS, these are site specific.

## Slide 135: Performance Test
Sampling Period
Must be no less than 30% of the cycle time for batch sampling CEMS.
Cycle/Response Time
The response time of PM CEMS, which is equivalent to the cycle time, must be no longer than 15 minutes.
In addition, the delay between the end of the sampling time and reporting of the sample analysis must be no greater than 3 minutes. Must document any changes in the response time following installation.
3 min.
15 min.

## Slide 136: Developing a PS-11 Correlation
Minimum of 15 valid RM test runs
Simultaneous PM CEMS data
Plot RM data vs. CEMS data
Does data meet the criteria of PS-11?

## Slide 137: Procedure 2 – Quality Assurance Requirements For Particulate Matter CEMS At Stationary Sources
What are the basic functions of Procedure 2?
Assess the quality of your PM CEMS data by estimating measurement accuracy
Control and improvement of the quality of your PM CEMS data by implementing QC requirements and corrective actions until the data quality is acceptable
Specify the requirements for daily instrument zero and upscale drift checks and daily sample volume checks, as well as routine response correlation audits (RCA), absolute correlation audits, sample volume audits (SVA), and relative response audits (RRA)
NOTE: Requires periodic evaluations of PM CEMS performance and the development and implementation of QA/QC programs to ensure that PM CEMS data quality is maintained.

## Slide 138: Procedure 2 – Quality Control Program
Your QC program must, at a minimum, include step-by-step procedures for the following:
    Performing drift checks
    Methods for making adjustments to PM CEMS
    Preventative maintenance
    Data recording, calculations, and reporting
    Performing RCA and RRA procedures
    Performing absolute correlation audits (ACA) and SVA
    Corrective actions for malfunctioning PM CEMS, including flagged data periods
    Procedures for checking extractive system ducts for material accumulation

## Slide 139: Procedure 2 – Auditing Requirements
Daily Checks | Quarterly Checks (performed 3 out of 4 quarters annually) | Annual Checks (frequency specified in permit or applicable regulation)
Zero and upscale drift | ACA | RRA – usually, annually, unless a RCA is performed
Check the system optics (light-scatter and extinction-type) | SVA | RCA – usually, once every three years
Sample volume check (if used in calculating output) |  | 

## Slide 140: Procedure 2 – Performance Criteria
Zero Or Upscale Drift | Sample Volume Measurement | Absolute Correlation Audits | Sample Volume Audits
Must be less than 4% for 5 consecutive day, or 8% for any one day | Must be less than 10% for 5 consecutive day, or 20% for any one day | Cannot exceed +10% of average audit value or 7.5% of applicable standard | Must be less than +5% of sample volume audit value

## Slide 141: Procedure 2 – Performing Response Correlation Audit Performance Criteria

## Slide 142: Procedure 2 – Relative Response Audit Performance Criteria
NOTE: If your PM CEMS fails to meet these RRA criteria, it is out of control.

## Slide 143: Commonly Used Technologies for PS-11
A light scattering PM CEMS measures the light scattered by the entrained particulate in the stack exhaust, the amount of scattering being proportional to the particulate concentration, and affected by particle size, shape, and color.
A beta gauge PM CEMS uses a beta radiation source and an adhesive filter tape material which collects the PM material at predetermined intervals. The collected PM on the filter tape attenuates the beta radiation, the amount of attenuation being proportional to the mass of collected PM, and independent of particle characteristics.
NOTE: Both PM CEMS require site-specific correlation against manual gravimetric RM measurements.
Images reprinted with permission from Thermo Fisher Scientific™

## Slide 144: 12. Let’s Test Your Knowledge!
144

## Slide 145: PS-11 – Question 1, 2
1. PS-11 is primarily used for certifying _____.
Check Answer
1. PS-11 is primarily used for certifying PM CEMS.
2. How does PS-11 differ from other performance specifications?
A. Type and characteristics of PM vary from source to source
B. Gas is easier to detect
C. Gas takes up a larger space, so it’s easier to capture
D. Analyzers might pick up dust in the air
2. A. Type and characteristics of PM vary from source to source.

## Slide 146: PS-11 – Question 3, 4
3. What are two commonly used measurement techniques for PM CEMS?
Check Answer
3. Light scattering and beta gauge are two commonly used measurement techniques for PM CEMS.
4. Which procedure is associated with PS-11?
Check Answer
4. Procedure 2 is associated with PS-11.

## Slide 147: Performance Specification 12A (PS-12A) Specifications and Test Procedures for Total Vapor Phase Mercury (Hg) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
147

## Slide 148: PS-12A Overview
Evaluates the acceptability of total vapor phase (gas-phase elemental and oxidized) Hg CEMS installed at stationary sources at the time of or soon after installation and whenever specified in the regulations.
    Regardless of whether it addresses elemental or oxidized Hg, the CEMS must record concentrations at standard conditions on a wet or dry basis
The Hg CEMS must be capable of measuring the total concentration in μg/m3 of vapor phase Hg, regardless of speciation, and recording that concentration at standard conditions on a wet or dry basis.
    These specifications do not address measurement of particle bound Hg.

## Slide 149: PS-12A Overview (Cont’d)
CEMS must meet the specified ranges:
  Zero-level gas - 0 to 20% of the span value
  Mid-level gas - 50 to 60% of the span value
  High-level gas - 80 to 100% of the span value

## Slide 150: Measurement Error and Calibration Drift
NOTE: Must perform ME, CD, RA, and linearity.
Measurement Error Test | Calibration Drift Test
For Hgo, the measurement error (ME) <5% of the span value at the zero-, mid-, and high-level reference gas concentrations. | CEMS <5% of the span value on any of the 7 days of the CD test. Use zero-level gas and either mid- or high-level gas.
For HgCl2, the ME <10% of the span value at the zero-, mid-, and high-level reference gas concentrations. | 

## Slide 151: Relative Accuracy and Reference Methods
Relative Accuracy Test | Reference Methods
CEMS <20% of the mean value of the RM test data in terms of units of μg/scm Alternatively, if the mean RM <5.0 μg/scm, the results are acceptable if the absolute value of the difference between the mean RM and CEMS values <1.0 μg/scm. | Method 29 and ASTM 6784, filterable portion not included. Determine number of sampling points by Method 1. Minimum of nine 2-hour test runs.
 | Method 30A and 30B Use 12 sampling points according to Method 1. Minimum of nine 30-minute test runs.

## Slide 152: Procedure 5 - Quality Assurance Requirements For Vapor Phase Mercury CEMS And Sorbent Trap Monitoring Systems Used For Compliance Determination At Stationary Sources
What are the basic functions of Procedure 5?
  To ensure Hg CEMS (vapor recovery or sorbent trap) meet acceptable standards for determining compliance on an ongoing basis:
        Assessment of the quality of Hg CEMS data
        Control and improvement of the quality of Hg CEMS data by implementing QC requirements and corrective actions
        Specification of QC requirements

## Slide 153: Procedure 5: Quality Control Requirements
Minimum Requirements
  CD checks
  CD determination & adjustment
  Weekly system integrity check procedures
  Routine operation, maintenance & QA/QC procedures for the sorbent trap monitoring systems
  Routine & preventative maintenance procedures (including spare parts inventory)
  Data recording, calculations & reporting
  Accuracy audit procedures
  Program of corrective action for malfunctioning CEMS

## Slide 154: Procedure 5: Calibration Drift Assessment

## Slide 155: Procedure 5: Data Accuracy Assessment
Hg CEMS Audit Requirements - an accuracy audit must be performed at least once each calendar quarter; successive quarterly audits (if possible) must be performed no less than two months apart
Relative Accuracy Test Audit | Alternative Quarterly Audits | Sorbent Trap Monitoring System Audit Requirements -
At least once every 4 calendar quarters except as noted in Section 5.1.4 of Appendix B; follow Section 8.5 of PS- 12A & calculate results according to Section 12.4 of PS-12A | Quarterly Gas Audit (QGA) - may be conducted in 3 of 4 calendar quarters but no more than 3 quarters in succession; challenge the CEMS with a zero and 2 upscale level audit gases of known concentrations, (20-30% of span & 50-60% of span) first of elemental Hg and then of oxidized Hg RAA - alternative to QGA; follow section 8.5 of PS-12A, but only 3 test runs required | RATA conducted at least once every 4 calendar quarters; perform the RATA as described in section 8.3 of PS-12B & calculate results per Section 12.4 in PS- 12A

## Slide 156: Procedure 5: Excessive Audit Inaccuracy
Out-of-Control Period Definition | Monitoring Data Status During Out-of-Control Period | Criteria for Excessive Audit Inaccuracy | Criteria for Acceptable QC Procedures
Hour immediately following the completion of failed RATA, RAA, or QGA or system integrity check until completion of subsequent successful test of the same type. | Cannot be used to determine compliance with an applicable emission limit or to meet minimum data availability requirements. | RATA: PS 12A& PS 12B – 20% or mean RM < 5.0 µg/scm if difference between CEMS and RM < 1.0 µg/scm QGA: +/- 15% of the average audit value or +/- 5 µg/m3 (whichever is greater) RAA: +/- 20% of the 3 run average or +/- 10% of the applicable standard, whichever is greater. | After 2 consecutive quarters with excessive inaccuracies, the owner/operator must revise the QC procedures or modify/repair/replace the CEMS which will require recertification of the CEMS.

## Slide 157: Commonly Used Technology for PS-12A
Hg sample is extracted from the stack and analyzed using atomic fluorescence spectroscopy to measure the concentration of Hg vapor in the sample.
  When an Hg atom absorbs the energy from a specific UV wavelength, an electron transitions from a stable ground state to an unstable, excited state, and when the UV energy source is removed, the electron returns to its stable state and emits a photon of light.

## Slide 158: Hg CEMS Atomic Fluorescence
Measure gaseous Hg
  Elemental (Hg0)
  Oxidized (Hg2+)
Almost all convert oxidized Hg to elemental Hg for measurement of total gaseous Hg
Calibrate using NIST-traceable Hg gas generators or cylinders

## Slide 159: 13. Let’s Test Your Knowledge!
159

## Slide 160: PS-12A – Question 1, 2
1. PS-12A is primarily used for certifying _____.
Check Answer
1. PS-12A is primarily used for certifying total vapor phase (gas-phase elemental and oxidized) Hg CEMS.
2. The results of a RATA test are considered valid if the calculated RA is less than ____ , or if the absolute difference is less than or equal to ____ ug/m3.
A. 20% and 1.0 ug/m3
B. 15.0% and 5.0 ug/m3
C. 20% and 5.0 ug/m3
D. 15.0% and 1.0 ug/m3
2. A. The results of a RATA test are considered valid if the calculated RA is less than 20%, or if the absolute difference is less than or equal to 1.0% ug/m3.

## Slide 161: PS-12A – Question 3, 4
3. What reference methods are commonly associated with PS-12A?
Check Answer
3. Reference methods M29, 30A, 30B, and ASTM 6784 are commonly associated with PS-12A.
4. What are two measurement techniques used for PS-12A?
Check Answer
4. Sorbent traps and atomic fluorescent spectroscopy are two measurement techniques used for PS-12A.

## Slide 162: Performance Specification 12B (PS-12B) Specifications and Test Procedures for Total Vapor Phase Mercury (Hg) Using Sorbent Traps in Stationary Sources
162

## Slide 163: PS-12B Overview 1
Establishes performance benchmarks for, and to evaluate the acceptability of, sorbent trap monitoring systems used to monitor total vapor-phase (gas-phase elemental and oxidized) Hg emissions in stationary source flue gas streams.
These monitoring systems involve continuous repetitive in-stack sampling using paired sorbent media traps with periodic analysis of the time-integrated samples.
The Hg monitoring system must be capable of measuring the total concentration of vapor phase Hg (regardless of speciation), in units of μg/dscm.
These procedures are only intended for use under relatively low particulate conditions (e.g., monitoring after all pollution control devices).

## Slide 164: PS-12B Overview 2
Known volumes of flue gas are continuously extracted through paired, in-stack, pre-spiked sorbent media traps at appropriate nominal flow rates.
Must use expected Hg concentration to determine sample flow rate and sorbent tube spike mass.
The sorbent traps in the sampling system are periodically exchanged with new ones, prepared for analysis as needed, and analyzed by any technique that can meet the performance criteria.
For QA purposes, a section of each sorbent trap is spiked with HgOx prior to sampling.
Following sampling, this section is analyzed separately, and a specified minimum percentage of the spike must be recovered. Paired train sampling is required to determine method precision.

## Slide 165: PS-12B Relative Accuracy and Reference Methods
For initial certification of a sorbent trap monitoring system, an RA test is required.
Sorbent trap used in RA must be same type material as used in daily operation.
Use 12 sampling points according to method 1.
The RA of the sorbent trap monitoring system must be <20% of the mean value of the RM test data in terms of units of μg/scm.
  Alternatively, if the RM concentration is <5.0 μg/scm, then the RA results are acceptable if the absolute difference between the means of the RM and sorbent trap monitoring system values <1.0 μg/scm.

## Slide 166: Quality Assurance/Quality Control Criteria for Sorbent Trap Monitoring Systems
Pre-test Leak Check | Post-test Leak Check | Ratio Of Stack Gas Flow Rate To Sample Flow Rate
< 4% of target sampling rate | < 4% of average sampling rate | <5% of the hourly ratios or 5 hourly ratios (whichever is less restrictive) may deviate from the reference ratio by more than ±25%.

## Slide 167: Relative Deviation, Quality Assurance/Quality Control
Paired sorbent trap agreement:
  ≤10% Relative Deviation (RD) if the average concentration is > 1.0 μg/m3
  ≤20% RD if the average concentration is ≤1.0 μg/m3 or if absolute difference between concentrations from paired traps is ≤ 0.03 μg/m3.

## Slide 168: Quality Assurance/Quality Control
  Spike recovery study: average recovery between 85% and 115% for each of the three spike concentration levels.
  Multipoint analyzer calibration: each analyzer reading within ±10% of true value and r2≥ 0.99.
  Analysis of independent calibration standard: within ±10% of true value.
  Spike recovery from section 3 of both sorbent traps: 75-125% of spike amount.

## Slide 169: Quality Assurance/Quality Control (Cont’d)
To validate sorbent trap monitoring system data, the acceptance criteria for the following five QC specifications must be met for both traps:
Post-monitoring leak check
Ratio of stack gas flow rate to sample flow rate
Paired trap agreement
Section 2 breakthrough
Section 3 spike recovery
NOTE: To validate an RA test run, both traps must meet the acceptance criteria for all five QC specifications.
Must perform ongoing QA according to requirements of 40 CFR 60, Appendix F, Procedure 5 (See slides under PS-12A for more information).

## Slide 170: Common Technology
Stack exhaust is sampled through a sorbent trap system which collects the gaseous elemental and oxidized Hg on the sorbent media.
Sorbent traps are sent to the lab, where the Hg sample is extracted from the stack and analyzed using atomic fluorescence spectroscopy to measure the concentration of Hg vapor in the sample.
  When a Hg atom absorbs the energy from a specific UV wavelength, an electron transitions from a stable ground state to an unstable, excited state, and when the UV energy source is removed, the electron returns to its stable state and emits a photon of light.

## Slide 171: Sorbent Trap Monitoring System Background
• Integrated sample measures total
gaseous Hg
• For post-PM control locations
• Paired traps, in-stack with 3 sections
• Proportional sampling

## Slide 172: 14. Let’s Test Your Knowledge!
172

## Slide 173: PS-12B – Question 1, 2, 3
1. PS-12A and PS-12B differ because PS-12A is targeted specifically towards sorbent trap testing.
True
False
1. False.
2. Pre- and post- leak checks are required for PS-12B sampling.
True
False
2. True.
3. A paired set of traps are used in the methodology of PS-12B.
True
False
3. True.

## Slide 174: PS-12B – Question 4
4. What is the spike recovery requirement from section 3 of sorbent traps?
A. 75-125% of spike amount
B. 90-110% of spike amount
C. 85-115% of spike amount
D. 80-120% of spike amount
4. A. 75-125% of spike amount

## Slide 175: Performance Specification 15 (PS-15) Specifications and Test Procedures for Extractive Fourier Transform Infrared (FTIR) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
175

## Slide 176: PS-15 Overview

## Slide 177: What is Fourier Transform Infrared ?
FTIR is an analytical technique used to obtain an IR spectrum of absorption or emission of a gas.
An FTIR spectrometer simultaneously collects high-spectral-resolution data over a wide spectral range. A fourier transform (a mathematical process) is required to convert the raw data into the actual spectrum which is compared to a library of spectra to find a match.
Image reprinted with permission from Thermo Fisher Scientific™

## Slide 178: Fourier Transform Infrared System Requirements
FTIR CEMS must be equipped with reference spectra bracketing the range of path length-concentrations (absorbance intensities) to be measured for each analyte.
The optical configuration of the FTIR system must be such that maximum absorbance of any target analyte is no greater than 1.0.
Additionally, the minimum absorbance of any target analyte must be at least 10 times the root mean square deviation (RMSD) noise in the analytical region.
Analytical package must:
Include data stored to write-protected medium
Store one interferogram per hour
Include all absorbance spectra, as well as all background spectra and interferograms
Include all calibration transfer standard (CTS) spectra and interferograms

## Slide 179: Quality Assurance/Quality Control
Periodic Quarterly or Semiannual QA/QC Checks include:
  Audit Sample- Unknown target analyte(s) analyzed by a CEMS operator.
  Audit Spectra- Analytical results must be within ±5% of the certified audit concentration for each analyte (plus the uncertainty in the audit concentration).
    Only tests the analytical program of FTIR CEMS
Independent Analysis of Spectra by EPA
  Submit three representative absorbance spectra
  Corresponding CTS spectra
  Corresponding background spectra and interferograms
  Spectra of associated spiked samples
  Analytical results for the sample spectra

## Slide 180: Calibration and Standardization
  Calibration and standardization includes:

## Slide 181: Analytical Procedure
Initial Certification - Perform evaluation procedures in section 6.0 FTIR Protocol using either:
  Method 301 validation procedures; or
  Comparison to applicable RM
Validation - Use EPA method 301. Procedures include spiking known concentrations of analytes and tracer gas (SF6) while sampling source gas. 2 options include:
  FTIR CEMS analyzing spectra collected sequentially (Validation run consists of 24 independent results- 12 spiked and 12 unspiked)
  FTIR CEMS operating side by side (Validation run consists of 24 independent results- 12 spiked and 12 unspiked)
Compare to a RM. Perform 9 runs of at least 30 minutes consisting of at least 5 independent FTIR CEM samples.

## Slide 182: Analytical Procedure – 3 Methods
    NOTE: For all three methods, use equations in PS-2 for RA determinations.

## Slide 183: Absorbance Spectrum Measured by Fourier Transform Infrared
Image reprinted with permission from Thermo Fisher Scientific™
https://www.thermofisher.com/us/en/home/industrial/spectroscopy-elemental-isotope-analysis/spectroscopy-elemental-isotope-analysis-learning-center/molecular-spectroscopy-information/ftir-information/ftir-basics.html

## Slide 184: 15. Let’s Test Your Knowledge!
184

## Slide 185: PS-15
1. What does the CEMS measure that PS-15 is used to certify?
Check Answer
1. The CEMS measures all HAPs, as well as volatile organic and inorganic species.
2. Which EPA FTIR reference method is associated with PS-15?
A. 320
B. 301
C. 365
D. 196
2. B. EPA FTIR reference method 301 is associated with PS-15.
3. Which technique is used by PS-15?
A. Integrated sampling
B. FTIR
C. Grab Sampling
D. CEMS as RM
3. B. FTIR is a technique used by PS-15.

## Slide 186: Performance Specification 16 (PS-16) Specifications and Test Procedures for Predictive Emission Monitoring Systems (PEMS) in Stationary Sources
186

## Slide 187: What is a PEMS?
PEMS refers to all the equipment that is required to predict an emission concentration or emission rate.
Unlike a CEMS which uses sampling and analytical equipment to directly measure specific pollutant concentrations, a PEMS uses the continuous measurement of selected plant parameters and plant operating conditions with a software-based system of mathematical models to determine the pollutant emissions.

## Slide 188: PS-16 Overview

## Slide 189: PS-16 Overview (Cont’d)

## Slide 190: Initial Certification
RA must be <10% if PEMS measures >100 ppm or 0.2 lbs/mmBtu.
RA must be <20% if PEMS measures <100 ppm (0.2 lb/mmbtu) and >10 ppm (0.02 lb/mmbtu).
RA if PEMS measures < 10 ppm, the absolute mean difference between the PEMS measurements and the RM measurements must not exceed 2 ppm (0.01 lb/mmbtu).
For diluent PEMS, an alternative criterion of ±1% absolute difference between the PEMS and RM may be used if less stringent.

## Slide 191: Initial Certification (Cont’d)
Must be performed at 3 load levels:
    Low load (between minimum stable load and 50%)
    Mid load (50 to 80%)
    High load (80 to 100%)
Bias Correction: If average difference < absolute value of confidence coefficient, no correction factor is needed.
PEMS Training: If Fcritical ≥ Fr ≥ 0.8, optional after initial and subsequent RATAs.
Annual RATA testing must be performed at normal load.

## Slide 192: Ongoing Quality Assurance Tests
Quarterly RAA
  May use portable analyzer (must meet ASTM D6522-00) or RM testing
  Three 30-minute test runs
First year, a RAA must be performed in 3 of the 4 quarters.
  If all three pass, and the 2nd year RATA passes, then only a semi-annual RAA may be required.
  If, at anytime, either a RAA or RATA test fails then quarterly RAAs must resume.

## Slide 193: Ongoing Quality Assurance Tests (Cont’d)
Test | PEMS Regulatory Purpose | Acceptability | Frequency
Sensor Evaluation | All |  | Daily
RAA | All | 3-test avg <10% of Simultaneous analyzer or RM average | Each quarter except quarter when RATA performed
Bias Correction | All | If davg ≥ ¦cc¦ | Bias test passed (no correction factor needed)
PEMS Training | All | If Fcritical ≥Fr≥0.8 | Optional after initial and subsequent RATAs
Sensor Evaluation Alert Test (optional) | All | See Section 6.1.8 | After each PEMS training

## Slide 194: 16. Let’s Test Your Knowledge!
194

## Slide 195: PS-16 – Question 1, 2
1. How does PEMS work?
Check Answer
1. PEMS uses the continuous measurement of selected plant parameters and plant operating conditions with a software-based system of mathematical models to determine the pollutant emissions.
2. PS-16 is used for determining whether a PEMS is acceptable for use in demonstrating compliance with applicable requirements.
True
False
2. True. PS-16 is used for determining whether a PEMS is acceptable for use in demonstrating compliance with applicable requirements.

## Slide 196: PS-16 – Question 3
3. During initial testing, how many runs must be used for excess emissions monitor and compliance monitor?
A. Excess-9 runs; compliance- 27 runs
B. Excess-12 runs; compliance- 12 runs
C. Excess-10 runs; compliance- 30 runs
D. Excess-9 runs; compliance- 3 runs
3. A. Excess - 9 runs; Compliance- 27 runs

## Slide 197: Performance Specification 18 (PS-18) Specifications and Test Procedures for Gaseous Hydrogen Chloride (HCl) Continuous Emission Monitoring Systems (CEMS) in Stationary Sources
197

## Slide 198: PS-18 Overview
NOTE: Extended performance assessment requirements found in App. F, Procedure 6.

## Slide 199: Interference Test
  Interference response(s) must not be >2.5% of the calibration span or ±3.0% of the equivalent HCl concentration used for the interference test (whichever is less restrictive), or
  The sum of the interference response(s) does not exceed six times the level of detection (LOD) or 0.5 ppmv for a calibration span of 5 to 10 ppm, or 0.2 ppmv for a calibration span of less than 5 ppmv.

## Slide 200: Beam Intensity and Temperature Verification
Notes: Integrated Path (IP): type of in situ CEMS

## Slide 201: Pressure Verification and Level of Detection
Pressure Verification Procedure (IP-CEMS only)
The absolute relative difference between the measured value of stack pressure (MP) and the pressure value from the calibrated pressure reference device (VP) must be ≤5.0%, or
The absolute difference between MP and VP must be ≤0.12 kilopascals (0.5 inches of water column), whichever is less restrictive.
Level of Detection Determination
Must determine the minimum amount of HCl that can be detected above the background in gas matrix.
  Determined in a laboratory or by manufacturer.
  Must be less than 20% of applicable limit.

## Slide 202: Response Time and Measurement Error Tests
Response Time Test
Measurement Error Test
Must determine measurement error (ME), level of detection (LOD) and standard addition (SA) response times.
3 sets of data are used to determine mean upscale and downscale response times for each procedure.
Extractive CEMS ME Test
  Measure 3 upscale HCl reference gases 80-100% of span.
IP-CEMS ME Test
  Conduct 3-level system ME test by individually adding the known concentrations of HCl reference gases into a calibration cell of known volume, temperature, pressure and path length.
  The ME must be less than or equal to 5.0% of the span value at the low-, mid-, and high-level reference gas concentrations.

## Slide 203: PS-18 Calibration Drift and Relative Accuracy
7-Day Calibration Drift Test
Relative Accuracy
Must complete before RA tests.
  Determine magnitude of CD at 24-hour intervals for 7 consecutive operating days, not necessarily 7 calendar days.
Must be established against RM 26A, 320, 321, or ASTM D6348-12.
  Conduct diluent, moisture, and pollutant measurements simultaneously.
  Test at 12 points, 6 points, 3-point long line or, if a stratification test is passed, the 3-point short line.
  Conduct a minimum of 9 RM runs.
  RA must be < 20.0% of RM or < 15.0% of RM if average emission level is <75% of emission standard.

## Slide 204: Quality Assurance Procedure 6
What are the basic functions of Procedure 6?
To ensure HCl CEMS data meets acceptable standards for determining compliance on an ongoing basis:
  Assessment of the quality of your HCl CEMS data
  Control and improvement of the quality of your HCl CEMS data by implementing QC requirements and corrective actions
  Specification of QC requirements

## Slide 205: Procedure 6
Requires that the CEMS is audited to assess the data accuracy
Temperature and pressure measurement devices must be audited annually
RATA must be conducted once every four calendar quarters
If HCL concentration is < 20 % of the concentration equivalent to the applicable emission standard, you must perform a CGA or dynamic spiking at least one quarter. Otherwise RAA must be performed quarterly.

## Slide 206: Procedure 6 (Cont’d)
If CEMS is in-control for 8 consecutive quarters and emits < 75 % of the concentration equivalent to the applicable standard, the auditing procedures may be revised to use CGA, RAA or dynamic spike audit (DSA) for 7 quarters before performing a RATA.
At a reporting interval specified in the permit or applicable regulation; the results of the quarterly audits and annual accuracy audit results as well as the daily assessment results must be reported to the appropriate agency.

## Slide 207: Commonly Used Technologies for PS-18
FTIR—refer to details for PS-15
Tunable diode laser (TDL) spectroscopy uses the absorbance spectra of target gases and the ability to tune the laser to a specific absorbance wavelength of the gas to measure the gas concentration. It can achieve very low detection limits (ppb), and it is also possible to determine the temperature, pressure, velocity and mass flux of the gas being measured. It is sometimes used as the light source in cavity ring-down spectroscopy (CRDS) (see details regarding this technology on the next slide).

## Slide 208: Commonly Used Technologies for PS-18 (Cont’d)
In CRDS, the beam from a single-frequency laser diode tuned to the absorbance of the gas being measured enters a cavity defined at least two high reflectivity mirrors with a path length in kilometers, making it extremely sensitive to very low concentrations of the target gas.
When the laser is on, the cavity quickly fills with reflected laser light.
A photodetector senses the small amount of light leaking through one of the mirrors to produce a signal that is directly proportional to the intensity in the cavity.

## Slide 209: 17. Let’s Test Your Knowledge!
209

## Slide 210: PS-18
1. A CEMS that is certified using PS-18 measures what pollutant?
Check Answer
1. A CEMS that is certified using PS-18 measures HCl.
2. RA must be established with what reference methods?
Check Answer
2. RA must be established with reference methods
26A, 320, 321, or ASTM D6348-12.
3. What is the basic function of procedure 6?
Check Answer
3. The basic function of procedure 6 is to ensure HCl CEMS data meets acceptable standards for determining compliance on an ongoing basis.
4. What are the common technologies used for PS-18?
Check Answer
4. Common technologies used for PS-18 include a TDL, CRDS, and FTIR.

## Slide 211: Activity
Title: Which Performance Specification is it?
Purpose: To review and recall module content associated with a specific performance specification
Time: 45 minutes
            30 minutes in groups
            15 minutes group debrief
211

## Slide 212: Activity Debrief
212

## Slide 213: Module 3 Summary
S
Now that you have completed Module 3, you should be able to:
  Define key terms, such as calibration drift (CD), relative accuracy, span, etc.
  Compare performance specification (PS) and quality assurance (QA) procedures by pollutant, where relevant
  List relevant QA procedures by PS
  Provide examples of technologies that can be used for each PS
  Recognize why predictive emission monitoring systems (PEMS) performance specifications are different from others