## Slide 1: Module 4: Audits/Inspections and Enforcement

## Slide 2: Module 4 Outline

## Slide 3: Module 4 Learning Objectives
At the end of Module 4, learners will be able to:
  Distinguish the difference between performance audits and systems/field audits
  Explain the utility of performance audits and systems/field audits
  Describe the inspector’s role during an audit
  Describe the procedures necessary to use CMS data in determining compliance
  Assess daily, weekly, monthly, quarterly, and annual required preventative maintenance and QA requirements
  Distinguish between CMS as compliance method and CMS data as credible evidence

## Slide 4: Overview of Audits and Inspections
Performance Audit
A quantitative evaluation, which includes things such as*: Cylinder Gas Audit (CGA) Relative Accuracy Test Audit (RATA) Relative Accuracy Audit (RAA)
System/Field Audit
Qualitative evaluation involving an inspection
*These elements are covered in Module 3

## Slide 5: Performance and System Audits
EPA relies on a combination of performance and system/field auditing to verify overall data integrity.

## Slide 6: Performance Audits
Required by 40 CFR Part 60
Found in performance specifications and QA procedures
Include the daily, quarterly, and annual audit procedures
Audit results usually submitted to agency for review

## Slide 7: System/Field Audit or Inspection
May be conducted in conjunction with a performance audit such as a RATA or RCA.
Allows the observer to…
  Physically inspect the CEMS,
  Review the data collected, and
  Review maintenance logs, etc.
  Note: The audit procedures for Part 75 can be found here: https://www.epa.gov/airmarkets/field-audit-manual

## Slide 8: Inspector’s Role in System/Field Audits
Every inspector should check with their agency to see exactly what their policies and procedures are before conducting an audit.
Typically, the inspector’s role is not to provide technical advice or consulting on the operation of the monitoring equipment.
Usually a "hands off" approach is used when conducting the audit so that the inspector does not have any physical contact with the monitoring system hardware.

## Slide 9: Inspections: Before Going Onsite
Preparing for an inspection:
Review any monitoring plans or test protocols, quality assurance/quality control (QA/QC) manuals, RATA records, quarterly audit records, and quarterly emission report submittals.
Check data availability, amount and causes of downtime, significant maintenance and any reports of replacement of key components.
Make note of multiple failed QA tests, missing data, unusual data trends (inconsistent over time, or inconsistent with other, similar facilities), and calculation errors.
9

## Slide 10: Inspections: Before Observing a Performance Audit
Preparing to observe any performance audit that will be conducted during the site visit:
Review the results of the last audits, noting any issues
Review the necessary performance specifications and QA procedures
Remember these performance specifications and QA procedures were covered in greater detail in Module 3 of this training
10
10

## Slide 11: 1. Let’s Test Your Knowledge!
11

## Slide 12: Audits/Inspections – Question 1
The observer of an audit should move the CEMS analyzers in order to be able to verify their model and serial numbers.
True
False
It’s false. The auditor should not have any physical contact with the monitoring system hardware.

## Slide 13: Audits/Inspections – Question 2
From the list below, which activity should not be done to prepare for an onsite visit?
A. Review records (e.g., monitoring plans, RATA, etc)
B. Check data availability reports
C. Be prepared to provide technical advice or consulting on the operation of the monitors
D. Make note of any data errors or issues
E. Review previous audits
F. Review performance specifications and QA procedures, where relevant
C. An inspector’s role is NOT to provide technical advice or consulting.

## Slide 14: Audits/Inspections – Question 3
A systems/field audit should never be conducted during the same period a performance audit is being performed.
True
False
It’s false. A system audit can be done at the same time as a performance audit. In many cases, it is more efficient to do so.

## Slide 15: Inspections Process – At the Facility
Verify:
Calibration and audit gas used, and
General appearance of analyzers and sampling system
Check:
Check if any alarms or warnings are indicated on DAS screen or system panels
Review:
Additionally, review maintenance logs and verify regular maintenance (daily, weekly, monthly) activities.
  Compare with description of these activities in QA/QC manual
  Note any frequent or reoccurring problems

## Slide 16: Visual Inspection of the CEMS
If possible, do a visual inspection of the CEMS from the sample probe location on the stack or duct, following the sample line to the CEMS shelter, and continuing inside the shelter through the gas conditioning system (if source-level extractive) to the analyzers.
Does the system look to be well maintained?
Are there low spots in the sample line where moisture might collect and scrub out pollutants? If the facility experiences cold winters, are all parts of the sampling system heated?
Is the physical location of the CEMS probe reasonable to access for maintenance?

## Slide 17: Visual Inspection of the CEMS (Cont’d)
Inside the CEMS shelter, check the condition of the sample gas conditioning system for any condensed liquid in Teflon lines.
Where does the liquid drain?
Could it get blocked or freeze?
Are there signs of corrosion of valves and fittings?

## Slide 18: CEMS Shelter

## Slide 19: Calibration Gas Cylinders

## Slide 20: Pressure on Gas Cylinder

## Slide 21: CEMS Inspection
Probe
Umbilical
Conditioning
Dilution Air
Calibration Gas
Analyzers
DAS and Audits
Data and Records

## Slide 22: Measurement of Emissions
CEMS must be representative:
When the effluents from two or more emission points are combined before being released to the atmosphere, the owner or operator:
May install applicable continuous monitoring systems on each effluent or on the combined effluent, if subject to the same emission standards
Must install separate continuous monitoring systems on each effluent, if not subject to the same emissions standard.
22

## Slide 23: Measurement of Emissions (Cont’d)
When the effluent from one affected facility is released to the atmosphere through more than one point, the owner or operator shall install an applicable continuous monitoring system on each separate effluent unless the installation of fewer systems is approved by the administrator.
Results must be reported for each CEMS.

## Slide 24: Who, What, Where, and Why
Importance of a pre-test meeting
  Who should be at the test...
  What will happen...
  Where is the meeting to be held...
  Why is this meeting necessary, etc....

## Slide 25: The Road to Continuous Compliance: Operation of CEMS
Conduct Calibration Drift (CD) at least once per operating day: Zero and Span - New Source Performance Standards (NSPS)
Must be in continuous operation. Except for system breakdowns, repairs, calibration checks, and zero and span adjustments. A “continuous operation” means a minimum of one cycle of operation (sampling, analyzing, and data recording) is completed or each successive 15-minute period.
Maintain and operate the CEMS in a manner consistent with good air pollution control practices and manufacturer’s written specifications.
Ensure the visual display or indication of operation is readily available on-site

## Slide 26: Daily Data Readings

## Slide 27: Road to Continuous Compliance Data Reduction
Data from CEMS shall be reduced to 1-hour averages of valid data as follows*:
Operating Hour is a: | Valid Data Means: | What is excluded?
Full Hour (60 minutes of operation) | At least four data points – one data point in each 15-minutes quadrant of the hour | Unavoidable CMS breakdowns, out-of-control periods, repairs, maintenance periods, calibration checks, and zero (low-level) and high-level adjustments
Partial Hour (less than 60 minutes of operation) | At least one data point in each 15-minutes quadrant of the hour of operation | Same
Operating Hour with Maintenance or QA and the CMS operates: | Two or more quadrants of the hour - a minimum of two data points, separated by at least 15 minutes; One quadrant of the hour - at least one valid data point (not applicable for Part 63) | Same
*Note: the information presented here is aligned with the General Provisions and may vary by Subpart.

## Slide 28: 1-Hour Average Validation
….1-hour averages shall be computed as follows…for a full operating hour (any clock hour with 60 minutes of unit operation), at least four valid data points are required to calculate the hourly average, i.e., one data point in each of the 15-minute quadrants of the hour. [§ 60.13 (h)(1) & (h)(2)(i)]
0

## Slide 29: Hourly Validation Example
Minute | Value
0 | 3
1 | 4
2 | Invalid
3 | Invalid
4 | Invalid
5 | Invalid
6 | Invalid
7 | Invalid
8 | Invalid
9 | Invalid
10 | Invalid
11 | Invalid
12 | Invalid
13 | Invalid
14 | Invalid
Minute | Value
15 | Invalid
16 | Invalid
17 | Invalid
18 | Invalid
19 | Invalid
20 | Invalid
21 | Invalid
22 | Invalid
23 | 4
24 | 4
25 | 3
26 | 3
27 | 3
28 | 2
29 | 2
Minute | Value
30 | 2
31 | 3
32 | 4
33 | 2
34 | 3
35 | 4
36 | 4
37 | 5
38 | 5
39 | Invalid
40 | Invalid
41 | Invalid
42 | Invalid
43 | Invalid
44 | Invalid
Minute | Value
45 | 2
46 | 4
47 | 5
48 | Invalid
49 | Invalid
50 | Invalid
51 | 5
52 | 4
53 | 2
54 | Invalid
55 | Invalid
56 | Invalid
57 | Invalid
58 | Invalid
59 | Invalid
 | No.
Valid | 2
Invalid | 13
Quadrant | Valid
 | No.
Valid | 7
Invalid | 8
Quadrant | Valid
 | No.
Valid | 9
Invalid | 6
Quadrant | Valid
 | No.
Valid | 6
Invalid | 9
Quadrant | Valid
0

## Slide 30: Preventative Maintenance Examples Handout
Refer to the handout, “Daily, Weekly, Monthly, Quarterly, and Annual Preventative Maintenance Examples” in your participant guide.

## Slide 31: 2. Let’s Test Your Knowledge!
31

## Slide 32: Audits/Inspections – Question 4
How should a CEMS be sited?
Check Answer
CEMS must be installed so that representative measurements of emissions are obtained.

## Slide 33: Audits/Inspections – Question 5
How is data from a CEMS reduced to 1-hour averages?
Check Answer
For a full operating hour, at least four valid points are required to calculate the hourly average (i.e., one data point in each of the 15-minute quadrants of the hour).

## Slide 34: Audits/Inspections – Question 6
What are some examples of parameters that an observer should be looking at for CEMS daily operations?
Check Answer
Sample flow, gas pressure, and moisture (there are certainly others)

## Slide 35: Use of CMS Data for Enforcement
On technical grounds, CMS data typically are at least comparable to compliance method and inspection data derived from equally well-executed and quality-assured monitoring.
CMS data are more representative of actual continuous emissions than are some traditional sources of compliance data, such as emission factors and engineering calculations.

## Slide 36: Use of CMS Data for Enforcement (Cont’d)
CMS is important to enforcement, irrespective of whether the legal requirement being enforced specifies CMS as the compliance method.
However, a governing regulation (e.g., 40 CFR Part 60) must specify CMS as the compliance method in order for EPA to rely on CMS data alone to prove a violation of an emission limitation in Federal district court, or to issue a Notice of Noncompliance ("NON").
The same is true if EPA is to rely on CMS data alone to issue an administrative order respecting emissions violations.

## Slide 37: Enforcement Applications of CMS
*e.g., 40 CFR Part 60
The next few slides will walk through these two enforcement applications.
The governing regulation* specifies CMS as the compliance method
The governing regulation* specifies some method other than CMS as the compliance method, or, the governing regulation doesn’t specify a compliance method.

## Slide 38: CMS is the Compliance Method
Required by some NSPS, National Emission Standards for Hazardous Air Pollutants (NESHAPS), and State Implementation Plans (SIPs)
Includes data validation requirements
Requires monitoring against emission limits with long averaging time
Data documents compliance against the emissions standard in the units of the emissions standard

## Slide 39: CMS is Not the Compliance Method
CMS data is “Credible Evidence”
Data is used for initiating and supporting enforcement cases alleging emissions violations.
CMS data may provide a basis to issue a section 114 request for compliance method data.
CMS data may be used to enforce operation and maintenance, monitoring and recordkeeping and reporting requirements, when the regulation does not specify a compliance method or an emissions standard (e.g. general duty clause).

## Slide 40: 3. Let’s Test Your Knowledge!
40

## Slide 41: Audits/Inspections – Question 7
How can CMS be used to determine compliance with an emissions standard?
Check Answer
CMS may be the compliance method specified in the rule or CMS may be used as credible evidence.

## Slide 42: Activity
Title: Geometric Close
Purpose: To review the module content by enabling participants to summarize what they learned.
Time: 40 minutes
            20 minutes in groups
            20 minutes group debrief

## Slide 43: Activity Debrief
43

## Slide 44: Module 4 Summary
Now that you have completed Module 4, you should be able to:
  Distinguish the difference between performance audits and systems/field audits
  Explain the utility of performance audits and systems/field audits
  Describe the inspector’s role during an audit
  Describe the procedures necessary to use CMS data in determining compliance
  Assess daily, weekly, monthly, quarterly, and annual required preventative maintenance and QA requirements
  Distinguish between CMS as compliance method and CMS data as credible evidence
S