UNITED STATES

ENVIRONMENTAL

PROTECTION AGENCY

# CONTINUOUS

# MONITORING SYSTEMS COURSE

PARTICIPANT gUIDE

TABLE OF CONTENTS

Welcome!	1

About this Course	1

Course Description	1

Course Goal(s)	1

Course Organization	1

Target Audience	2

Prerequisites	2

Evaluation	2

Course Agenda	3

About the Participant Guide	5

PowerPoint Slides	5

Appendices	5

Course Introduction	7

Module 1: Introduction to CMS	15

Module 2: Overview of CMS and CMS Design and Components	42

Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies	71

Module 4: Audits/Inspections and Enforcement	180

# Welcome!

## About this Course

### Course Description

This three-day classroom Continuous Monitoring Systems (CMS) Course provides you with a basic understanding of CMS. It provides an overview of how to perform regulatory reviews and includes key concepts and information on CMS, including types, regulations, analytical techniques, systems design and components, performance specifications, quality assurance (QA) requirements, commonly used technologies, audits/inspections, and enforcement procedures. In addition to the instructor presentation and lecture, you will have opportunities to apply your knowledge through practical exercises, facilitated discussions, and knowledge check games.

### Course Goal(s)

The CMS course has been designed as a three-day classroom immersion into:

- Key concepts pertaining to CMS

- Terminology and techniques that are common in the field

- Performance specifications and QA procedures

- How to use the performance specifications and where to locate requirements

- Different types of audits and enforcement procedures

### Course Organization

There are four modules in this course and an introductory module:

- Course Introduction

- Module 1: Introduction to Continuous Monitoring Systems (CMS)

- Module 2: Overview of CMS and CMS Design and Components

- Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies

- Module 4: Audits/Inspection and Enforcement

### Target Audience

This course is intended for a fairly broad audience and could include participants with non-technical backgrounds (e.g., a basic degree in science) as well as technical backgrounds (e.g., a chemical engineering degree). No prior knowledge of CMS is presumed. This course is intended primarily for new hires or any participants who need an understanding of CMS in their job responsibilities.

Staff who could benefit from the course include, but are not limited to the following:

- Staff who are inspection and/or enforcement personnel

- Staff who observe stack test and CMS certifications

- Staff who are reviewing audit reports

- Staff who work on permits

### Prerequisites

There are no prerequisites for this course.

### Evaluation

The following types of evaluation will be conducted:

- Level 1 evaluation (reaction): accomplished through the use of the standard EPA end-of-course evaluation

- Level 2 evaluation (learning): accomplished through the completion of a course exam and optional activities conducted throughout the course

### Course Agenda

| Scheduled Time | Module | Duration |
| --- | --- | --- |
| Day 1 |  |  |
| 8:30 am – 9:00 am | Course Introduction | 30 minutes |
| 9:00 am – 10:00 am | Module 1: Introduction to Continuous Monitoring Systems (CMS) | 60 minutes |
| 10:00 am – 10:30 am | Break | 15 minutes |
| 10:30 am – 12:00 pm | Continue Module 1: Introduction to Continuous Monitoring Systems (CMS) | 90 minutes |
| 12:00 pm – 1:00 pm | Lunch | 60 minutes |
| 1:00 pm – 1:30 pm | Wrap-up Module 1: Overview of the Course | 30 minutes |
| 1:30 pm – 2:00 pm | Module 2: Overview of CMS and CMS Design and Components | 30 minutes |
| 2:00 pm – 2:15 pm | Break | 15 minutes |
| 2:15 pm – 4:00 pm | Wrap-up Module 2: Overview of CMS and CMS Design and Components | 105 minutes |
|  | Total: | 435 minutes (7.5 hours) |
| Day 2 |  |  |
| 8:30 am – 10:00 am | Begin Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies | 90 minutes |
| 10:00 am – 10:15 am | Break | 15 minutes |
| 10:15 am – 12:00 pm | Continue Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies | 105 minutes |
| 12:00 pm – 1:00 pm | Lunch | 60 minutes |
| 1:00 pm – 2:00 pm | Continue Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies | 60 minutes |
| 2:00 pm – 2:15 pm | Break | 15 minutes |
| 2:15 pm – 4:00 pm | Continue Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies | 105 minutes |
|  | Total: | 450 minutes (7.5 hours) |

| Day 3 |  |  |
| --- | --- | --- |
| 8:30 am – 9:45 am | Module 3 and Wrap Up: Performance Specifications, Quality Assurance, and Commonly Used Technologies | 75 minutes |
| 9:45 am –10:00 am | Break | 15 minutes |
| 10:00 am – 12:00 pm | Begin Module 4: Audits/Inspections and Enforcement | 120 minutes |
| 12:00 pm – 1:00 pm | Lunch | 60 minutes |
| 1:00 pm – 2:00 pm | Module 4: Audits/Inspections and Enforcement | 60 minutes |
| 2:00 pm – 3:00 pm | Course Wrap-up, Post-Training Assessment (optional) and Exam | 60 minutes |
|  | Total: | 390 minutes (6.5 hours) |

## About the Participant Guide

### PowerPoint Slides

This participant guide contains a copy of the PowerPoint presentation slides that the instructor(s) will use to teach this course. This includes PowerPoint slides for the course introduction and for modules 1-4. This participant guide is yours to keep, so feel free to take class notes on the lines provided below each slide.

### Appendices

There are three appendices associated with this course, which will be provided to you by your instructor:

- Appendix A: Master Glossary – The glossary encompasses a complete list of terms and definitions for all of modules in this course.

- Appendix B: Master Resources – A list of resources, including helpful website links, are provided per module.

- Appendix C: Pre and Post Self-Assessment – Your instructor may direct you to complete a self-assessment prior to and after the course. These self-assessments will help you to establish a baseline of what you may already know about the information contained in this course. Then, a post self-assessment will allow you to compare how your knowledge and skills have changed as a result of completing the course.

### Handouts

In addition to the appendices listed above, there are several handouts associated with this course, which will be provided to you by your instructor:

- General provisions

- Properties of light

- Performance testing

- Commonly used technologies

- Commonly used acronyms

- Performance specifications

- CMS definitions by performance specification (PS)

- Preventative maintenance examples

COURSE INTRODUCTION

# Course Introduction

Course Introduction Description:

The course introduction serves as an overview of the CMS course. This overview includes the course goal, as well as an outline of the modules covered in the course. The course introduction also gives you a chance to meet the instructors, as well as the other participants taking the course. Instructors also use this time to cover administrative aspects of the course including a review of the course materials and course agenda.

MODULE 1: INTRODUCTION TO CMS

# Module 1: Introduction to CMS

Module 1 Description:

In Module 1, you will gain a basic understanding of continuous monitoring systems (CMS), including learning about the four different types of CMS. You will learn the reasons why CMS are important, as well as their many uses. Additionally, you will learn about the inclusion of CMS into Federal regulations. This module will introduce performance specifications and quality assurance as it relates to continuous emission monitoring systems (CEMS) and the use of CEMS data for enforcement.

Module 1 Objectives:

- Recognize the different types and uses of CMS

- Identify the regulations that contain CMS requirements

- Recognize, in general, what performance specifications are and how they are used

- Recall enforcement aspects of CMS

MODULE 2: OVERVIEW OF CMS AND CMS DESIGN AND COMPONENTS

# Module 2: Overview of CMS and CMS Design and Components

Module 2 Description:

In Module 2, you will be provided an overview of continuous monitoring systems (CMS) and their components and design. You will learn about opacity and how continuous opacity monitoring systems (COMS) are used. In addition, you will learn about CMS pollutant parameters, extractive and in-situ systems, and CEMS location and siting considerations. This module will also describe how continuous emission rate monitoring systems (CERMS) function.

Module 2 Objectives:

- Define opacity and describe how COMS are used

- Recognize the pollutant parameters measured by CMS

- Distinguish between extractive and in-situ systems

- Describe how CERMS function

- Give examples of CEMS location and siting considerations

MODULE 3: PERFORMANCE SPECIFICATIONS, QUALITY ASSURANCE, AND COMMONLY USED TECHNOLOGIES

# Module 3: Performance Specifications, Quality Assurance, and Commonly Used Technologies

Module 3 Description:

In Module 3, participants will be provided an overview of performance specifications (PS) used for evaluating the acceptability of continuous monitoring systems (CMS) at the time of, or soon after, installation and wherever specified in the regulation. They will also learn about the associated quality assurance (QA) procedures, detailed information about each PS, and commonly used technologies.

Module 3 Objectives:

- Define key terms, such as calibration drift (CD), relative accuracy (RA), span value, etc.

- Compare performance specification (PS) and quality assurance (QA) procedures by pollutant, where relevant

- List relevant QA procedures by PS

- Provide examples of technologies that can be used for each PS

- Recognize why predictive emission monitoring systems (PEMS) PS are different from others

|  | scr |  |
| --- | --- | --- |

MODULE 4: AUDITS/INSPECTIONS AND ENFORCEMENT

# Module 4: Audits/Inspections and Enforcement

Module 4 Description:

In Module 4, you will be provided an overview of audits and inspections used for CMS. You will learn about the types of audits and what is involved in an inspection. In addition, you will learn about preventative maintenance and failure, as well as enforcement.

Module 4 Objectives:

- Distinguish the difference between performance audits and systems/field audits

- Explain the utility of performance audits and systems/field audits

- Describe the inspector’s role during an audit

- Describe the procedures necessary to use CMS data in determining compliance

- Assess daily, weekly, monthly, quarterly, and annual required preventative maintenance and QA requirements

- Distinguish between CMS as compliance method and CMS data as credible evidence