MODULE 2: OVERVIEW OF CMS AND CMS DESIGN AND COMPONENTS

# Module 2: Overview of CMS and CMS Design and Components

Module 2 Description:

In Module 2, participants will be provided an overview of continuous monitoring systems (CMS) and their components and design. They will learn about opacity and how continuous opacity monitoring systems (COMS) are used. In addition, they will learn about CMS pollutant parameters, extractive and in-situ systems, and CEMS location and siting considerations. This module will also describe how continuous emission rate monitoring systems (CERMS) function.

Module 2 Learning Objectives:

- Define opacity and describe how COMS are used

- Recognize the pollutant parameters measured by CMS

- Distinguish between extractive and in-situ systems

- Describe how (CERMS) function

- Give examples of CEMS location and siting considerations

Module 2 Preparation:

- Provide post-it notes to each participant for two knowledge check games and a culminating Module 2 group activity.

- Title a flip-chart page “Top Ten” for the culminating Module 2 group activity.

Estimated Time: 150 minutes (2.5 hours)

| Slide 1 | Module 2: Overview of CMS and CMS Design and Components Instructor Notes: DIRECT: - Before introducing Module 2, direct participants to the Light Properties handout to read and review. This will help give participants context to the idea of using light to measure pollutant concentrations, which is introduced on Slide 5. - Once participants have completed reading and reviewing the Light Properties handout, introduce Module 2. SAY: “Module 2 is intended as an overview of CMS and its design components. Let’s look more closely at the topics this module will cover.” |
| --- | --- |

| Slide 2 | Module 2 Outline Instructor Notes: - Review the topics covered in the Module 2 outline. |
| --- | --- |

| Slide 3 | Module 2 Learning Objectives Instructor Notes: SAY: “Here are the objectives that you should be able to accomplish after we complete this module.” - Read the Module 2 learning objectives. |
| --- | --- |

| Slide 4 | Types of CMS? Instructor Notes: - Read the four types of CMS. SAY: “This module will provide an overview of CEMS, COMS and CERMS. PEMS will be covered later in Module 3, when we discuss Performance Specification 16.” |
| --- | --- |

| Slide 5 NOTE: With the exception of gas velocity, all monitors measure in units of concentration (e.g. %, ppm, ppb, ug/m3 or mg/m3). Opacity is the measurement of the ability of the particulate matter in the emission stream to block the transmittance of light through it. Gas velocity is the average velocity of the exhaust flowrate at the measurement location. With the cross-sectional area, you can determine the volumetric flow rate. The flow rate AND the pollutant concentration can give you the mass emission rate of the pollutant. | Pollutant Parameters Instructor Notes: - Explain that CMS may be used to measure the pollutant parameters listed on this screen. - Explain that CEMS may contain a combination of the following components: - An opacity monitor – when combusting solid or liquid fuels, or processing solid materials (wood, stone, lime, etc.), that will result in some amount of particulate matter (PM) emissions. - One or more pollutant gas concentration monitors – depending on what pollutants are emitted by the source and regulated with an emission limit. - Diluent gas monitor – oxygen or carbon dioxide. - Stack velocity – to determine exhaust volumetric flowrate. IMPORTANT/KEY MESSAGE: - Highlight that the list on this slide shows the most common measured pollutant parameters. |
| --- | --- |

| Slide 6 | General Categories of CMS Instructor Notes: - Explain that this diagram lists the types of categories of CMS components, measurement technology, and basic design. - Explain that it also includes the type of parameter being measured (e.g. gas, PM, velocity) and the basic design (e.g. extractive, in situ, parametric). - Explain that pollution is either gaseous or non-gaseous (particulate, droplet, mist). IMPORTANT/KEY MESSAGE: - Highlight that several pollutant concentration analyzers use some kind of measurement technology based on how light interacts with the pollutant, although some use an electroanalytical measurement technology (not light-based). |
| --- | --- |

| Slide 7 | Opacity and COMS Instructor Notes: SAY: “Let’s discuss opacity and COMS first. In the next few slides we will begin by discussing opacity then cover opacity monitoring systems in the later slides.” |
| --- | --- |

| Slide 8 NOTE: Not all PM is visible. It may not be enough to be seen by naked eye, but still enough to exceed a low mass emission limit. PM may be any color including black, white, brown, red, etc. It may be unburned carbon from poor fuel combustion, dust from manufacturing process, or condensation of uncontrolled organic pollutant emissions. The plume opacity is determined at the exit of the stack (or after any condensation) using US EPA Method 9. | Opacity – Setting the Stage Instructor Notes: - Define opacity (Op) and transmittance (Tr) and how it relates to COMS. |
| --- | --- |

| Slide 9 NOTE: Scattering and absorption of light by a stack plume gives evidence to its opacity. Transmittance is determined from 2 principle measurements: - IO (Reference Measurement of Light Intensity) - I (Light Intensity measured after it passes through flue gas) Ratio of 2 Intensities I/ IO = Fractional Transmittance Opacity = 100 - Transmittance or % Transmittance = 100 - %Opacity | Transmittance vs. Opacity Instructor Notes: - Use the image to show the relationship between transmittance and opacity. - Explain that transmissometers measure the transmittance of light that passes through the flue gas. IMPORTANT/KEY MESSAGE: - Highlight that different units of measure of transmission and opacity show the effect of PM on light passing through it. |
| --- | --- |

| Slide 10 | Opacity Instructor Notes: - Review the characteristics of opacity. IMPORTANT/KEY MESSAGE: - Highlight that opacity is a “quick and dirty” and much cheaper method to gauge a source’s likely compliance status with PM limit. However, a source may be in compliance with opacity limit and still exceed an allowable PM limit. |
| --- | --- |

| Slide 11 | COMS Instructor Notes: - Review the features of COMS on the slide. IMPORTANT/KEY MESSAGE: - Highlight that COMS are an indicator of the amount of particulate matter, but do not measure it directly. |
| --- | --- |

| Slide 12 | CEMS Design Instructor Notes: SAY: “Next, we will discuss CEMS design.” |
| --- | --- |

| Slide 13 NOTE: Extractive: gas transport, can be diluted or not. Extractive systems are easy to construct and reliable. They are free of PM, H2O is removed, or diluted below dew point, and sample cooled. Hot-wet systems for hydrocarbon (HC) do not require removal of H2O or cooling. Equipment needed includes valves, pumps, chillers; gas transport (umbilical line), and gas conditioning. | Basics of CEMS Design Instructor Notes: - Explain that CEMS can be divided into two general categories based on the means by which the sample gas is acquired (captured) and delivered to the analyzer: extractive systems and in-situ systems. - Review each type of system |
| --- | --- |

| Slide 14 NOTE: Remove as much water and particulate as possible before it gets to the sampling pump. More components focus on sample gas conditioning. | General Extractive System – Conditioning Cabinet Instructor Notes: - Explain a general extractive system CEMS design. |
| --- | --- |

| Slide 15 | General Extractive System Components Instructor Notes: - Explain that the following components listed on this slide are part of the general extractive system. |
| --- | --- |

| Slide 16 NOTE: The source-level extractive sample line may consist of multiple lines. It might include lines for sample collection, calibration gas injection, purge/blowback, and a spare. | Source-Level Extractive System Instructor Notes: - Explain the source-level extractive system as it relates to CEMS design. - Use the image to review the components of the source-level extractive system including: - Opacity Monitor - Sampling Probe for Extractive System - Flow Monitor - Gas Conditioning System - Analyzers - Data Acquisition System - Highlight the following characteristics of source-level extractive systems: - The stack gas sample is withdrawn from the stack. - In a non-dilution extractive system, the sample is often transported through a heated sample line to the sample conditioning system or to the analyzer, whichever is needed, to prevent moisture condensation which could scrub certain water-soluble pollutants or even freeze in cold winter climates. - Conditioning of the sample gas is required to remove any moisture and remaining fine particulate to prevent damage to the gas concentration analyzers. IMPORTANT/KEY MESSAGE: - If the gas concentration is measured on a dry basis and the stack flowrate is measured on a wet basis, a measurement of actual stack gas moisture or an assumed moisture is necessary to calculate a mass rate measurement of the pollutant. |
| --- | --- |
| Slide 17 NOTE: Umbilical bundle may contain 3 or 4 lines. This may include one for the sample gas, calibration gas, and ambient air conditioning. The dilution system umbilical is similar to the source-level extractive systems, except that dilution extractive system has a dilution air line to the probe. | Dilution Extractive Systems Instructor Notes: - Explain the dilution extractive system as it relates to CEMS design. - Use the image to review the components of the dilution extractive system. |

| Slide 18 | Dilution Extractive Systems (Cont’d) Instructor Notes: - Discuss the characteristics of dilution extractive systems. - Discuss the difference between the two types: - In-stack (where the dilution of the flue gas is performed in the probe) - Out of stack (where the dilution componentry of the system is at the end of the sample probe immediately outside the stack) |
| --- | --- |

| Slide 19 | Dilution Probe Instructor Notes: - Explain the dilution probe as it relates to CEMS design. - Use the image to review the components of the dilution probe. |
| --- | --- |

| Slide 20 | Dilution Probe Orifice Instructor Notes: - Explain the dilution probe orifice as it relates to CEMS design. - Use the image to review the components of the dilution probe orifice. |
| --- | --- |

| Slide 21 | Reasons to Consider Using Dilution Probes Instructor Notes: - State the reasons to consider using dilution probes. |
| --- | --- |

| Slide 22 | Overview of Measurement Bias Instructor Notes: - Define measurement bias and explain how it relates to CEMS design. - Explain the difference between systematic error and random error. |
| --- | --- |

| Slide 23 | General Sources of Bias for Extractive Systems Instructor Notes: - Review the general sources of bias for extractive systems. - Explain that some of the effects of bias are a larger concern in dilution extractive, source level extractive, or both. - Explain the difference between absorption vs adsorption. |
| --- | --- |

| Slide 24 | Sources of Biases in Dilution Probes Instructor Notes: - Explain the sources of biases in dilution probes. |
| --- | --- |

| Slide 25 | Considerations for Extractive System Components Instructor Notes: - Review the components and their characteristics as it relates to extractive systems. |
| --- | --- |

| Slide 26 Note: Performance Specification 18 refers to in-situ path CEMS as “Integrated Path CEMS.” Path systems - effluent passing by a specific monitor’s “line of sight” (which may range from several feet to the full distance across the stack or duct diameter) is analyzed. | In-Situ Systems Instructor Notes: - Describe in-situ systems as it relates to CEMS design. - Explain the different between point systems and path systems. |
| --- | --- |

| Slide 27 | In-Situ Monitors Instructor Notes: - Explain in-situ monitors as it relates to CEMS design. - Use the image to review the components of in-situ monitors. |
| --- | --- |

| Slide 28 | Advantages of Extractive Systems Instructor Notes: - Explain the advantages of extractive vs. in-situ systems. |
| --- | --- |

| Slide 29 | Disadvantages of Extractive Systems Instructor Notes: - Explain the disadvantages of extractive vs. in-situ systems. |
| --- | --- |

| Slide 30 | Advantages of In-Situ Systems Instructor Notes: - Explain the advantages of an in-situ vs. extractive system. |
| --- | --- |

| Slide 31 | Disadvantages of In-Situ Systems Instructor Notes: - Explain the disadvantages of an in-situ vs. extractive system. - For “Does not allow for expansion,” explain that this means no additional gases or time-sharing. - For “Path type may not be able to be located downstream of sorbent injection or spray dryer systems,” explain that this is due to potential interference by particulate or droplets. |
| --- | --- |

| Slide 32 | CEMS Components Instructor Notes: SAY: “We’ll now discuss the components of CEMS.” |
| --- | --- |

| Slide 33 NOTE: To avoid plugging the sample probe when PM or moisture is present, a filter can be placed on the end. Also, a baffle can be attached to deflect PM from the filter surface A filter can be made of: - Sintered stainless steel (porous); or - Ceramic materials | External Probe Filters Instructor Notes: - Explain external probe filters as it relates to CEMS. - Use the image to review the components of external probe filters. - Explain that the graphics to the left show the front views of the probe tip and that the graphics in the middle show the side views. - Explain that the arrows show how particulate matter goes around the different types of filtering systems. |
| --- | --- |

| Slide 34 NOTE: The filter can become blocked - Material adhering to the filter reacts with the gas to be measured. In wet scrubbers: - Calcium or lime may attack filter surface - “Scrub” SO2 from the sample gas - Blow back systems may be helpful - Minimizing filter blockage - Blow back times vary from every 15 minutes to once per hour. Note that it could take a few minutes for the PM to blow out. This could produce “no data” for a while. - Several minutes in duration | Filter Blow Back Instructor Notes: - Explain the filter blow back as it relates to CEMS. - Use the image to review the components of a filter blow back. - Explain that the filter blow back shows particulate matter coming out of the probe tip. |
| --- | --- |

| Slide 35 NOTE: Some systems incorporate a coarse filter (10-50 m porosity) at the probe tip and an internal fine filter (1 m pore size) at the flange assembly. | Internal Coarse Filter Instructor Notes: - Explain the internal course filter as it relates to CEMS. - Use the image to review the components of an internal coarse filter. - Explain that this image shows an internal filter, shows the potential for a bias check, due to the way in which the calibration gas enters the probe tip and gets drawn down through the probe, back to the analyzers. |
| --- | --- |

| Slide 36 | Inertial Filter Instructor Notes: - Explain the inertial filter as it relates to CEMS. - Use the image to review the components of an inertial filter. - Explain that a major problem associated with previous probes is plugging the filters. - Explain that in inertial filer, the air moves very quickly, and a pump pulls the flue gas through a cylindrical filter. - Explain that the inertial filter material can be made of sintered stainless steel or porous ceramic with a pore size of 2-5 m. |
| --- | --- |

| Slide 37 | Example of How a Dilution Probe Works Instructor Notes: - Explain the dilution probe as it relates to CEMS. - Use the image to review the components of a dilution probe and how it works. |
| --- | --- |

| Slide 38 OPTIONAL KNOWELDGE CHECK ACTIVITY Conduct the knowledge check questions as a whole group eliciting answers from participants in the class. | - Let’s Test Your Knowledge! Instructor Notes: SAY: “Now, let’s participate in a game to test your knowledge of the information you’ve learned so far.” ACTIVITY - Divide participants into two groups. - Use the flipchart to capture team names and points. - Ensure each group has post-it notes. - Explain the directions to the participants: - Instructors will present the first knowledge check on the screen. Groups will quietly discuss their answer and put it on a post-it note. - The instructor will indicate when all groups should show their answer. - The instructor will record a point for each group/team that answers correctly. This game format continues for all the knowledge check questions. - Tell participants that this game will continue until the end of Module 2 since there are additional questions to come. Explain that those groups/teams who gets the most points will receive a reward (e.g., a candy treat). - Proceed to the following slides to conduct the activity. |
| --- | --- |

| Slide 39 | CEMS Design – Question 1 Instructor Notes: - __________ systems condition the sample gas before analysis. Answer: D. Source Level Extractive. |
| --- | --- |

| Slide 40 | CEMS Design – Question 2 Instructor Notes: - In-situ CEMS require __________ equipment for a single sample point compared to extractive CEMS. Answer: B. Less. |
| --- | --- |

| Slide 41 | CEMS Design – Question 3 Instructor Notes: - True or False? Transmittance is the percentage of light that is attenuated by an optical medium. Answer: False. Opacity is the percentage of light that is attenuated by an optical medium. |
| --- | --- |

| Slide 42 | CEMS that Include a Flow Monitor (Continuous Emission Rate Monitoring Systems or CERMS) Instructor Notes: SAY: “Lastly we will discuss CEMS that include a flow monitor. In other words, continous emission rate monitoring systems or CERMS.” - Explain that most CEMS systems can be classified as in-situ or extractive. - Explain that opacity and flow monitors are in-situ monitors where flue gas is monitored in-place. - Explain that a flow monitor is used in conjunction with gas concentration measurements to calculate mass emission rates. - State that most Part 75 Sources are required to have a flow monitor. |
| --- | --- |

| Slide 43 NOTE: Velocity measurements are needed to calculate a volumetric flow rate, which is used with the pollutant concentration converted to lb/ft3, to determine pollutant mass emission rate (PMER). PMER = Cs QsAs = (Conc.) (Velocity) (Area) | Continuous Emission Rate Monitoring Systems (CERMS) Instructor Notes: - Describe CERMS. - Explain that CERMS are used in conjunction with gas concentration measures to calculate mass emission rates. - State that CERMS are required for most sources subject to 40 CFR Part 75. Review PMER as it relates to CERMS. |
| --- | --- |

| Slide 44 NOTE: The volumetric flow rate can be determined by the measure of flue gas velocity and cross-sectional area of the stack. Stack gases approaching the impact face: - Stagnate and create pressure before diverting around the pitot tube - Pressure is measured by fluid in a Manometer - No gases are collected (sampled, etc.) by the pitot tube. | Differential Pressure Monitoring Instructor Notes: - Explain differential pressure measuring as it relates to CERMS. Explain in particular that it measures at a certain number of traverse points |
| --- | --- |

| Slide 45 | Ultrasonic Flowmeter Instructor Notes: - Explain the ultrasonic flowmeter as it relates to CEMS. - Use the image to review and explain the functionality of an ultrasonic flowmeter. |
| --- | --- |

| Slide 46 | Considerations When Choosing a Location for CEMS Instructor Notes: - Review the considerations listed on the slide, when choosing a location for a CEMS system. - Explain that prior to installing a CEM system: - Reference method sampling should be done to verify that measured emissions will be representative of actual emissions at all operating conditions, and - Pollutant concentration and stack flowrate sampling points will be at locations that provide a REPRESENTATIVE concentration and flowrate across the entire stack cross-sectional area. - Explain that this should be done for all measurements required of the CEM system: gas concentrations and stack velocity. In some cases, a system may have to sample simultaneously from several locations across the stack cross-section to get a representative sample or measurement. - Explain that the source chooses the location according to the applicable performance specification and as long as the CMS can meet the required performance and QA criteria then it is deemed acceptable. |
| --- | --- |

| Slide 47 | Access for Reference Method Testing Instructor Notes: - Explain that there should be room for maintenance and repair of monitoring equipment, and for reference method testing equipment and test team for certification testing. - State that the long straight stack run provides laminar flow and homogenized, well-mixed exhaust flow. - State that lights are used when stack testing inevitably goes into the night. - From a safety perspective, the cage is important |
| --- | --- |

| Slide 48 OPTIONAL KNOWLEDGE CHECK ACTIVITY Conduct the knowledge check questions as a whole group eliciting answers from participants in the class. | 2. Let’s Test Your Knowledge! Instructor Notes: SAY: “Now, let’s continue our game and continue testing more of your knowledge of Module 2.” ACTIVITY - Tell participants to divide back into their same two groups. - Use the flipchart to continue to record points. - If necessary, briefly explain the directions to the participants again. - Instructors will present the first knowledge check on the screen. Groups will quietly discuss their answer and put it on a post-it note. - The instructor will indicate when all groups should show their answer. - The instructor will record a point for each group/team that answers correctly. This game format continues for all the knowledge check questions. - Explain that those groups/teams who gets the most points will receive a reward (e.g., a candy treat). - Proceed to the following slides to conduct the activity. |
| --- | --- |

| Slide 49 | CERMS – Question 1 Instructor Notes: - True or False? For CERMs, the calculation of pollutant mass emission rate requires velocity measurements. Answer: True |
| --- | --- |

| Slide 50 | CERMS – Question 2 Instructor Notes: - A CERMS is used when which of the following is required: - The gas concentration - The mass emission rate - Both Answer: B. The mass emission rate. |
| --- | --- |

| Slide 51 | CEMS Instructor Notes: - Important factor(s) for installation of CEMS is/are: - Accessibility - Representativeness - Sufficient distance from flow disturbances - Protection from weather and vibration - All of the above Answer: E. All of the above. |
| --- | --- |

| Slide 52 Purpose: To review the module content by sharing a list of top ten things learned about CMS and its design and components. Estimated Time: 40 minutes - 20 minutes in groups - 20 minutes in group debrief Preparation Checklist: - Title a flip-chart page “Top Ten” and post in the front of the room. NOTE: While groups are presenting, instructors are encouraged to walk around and observe. | Optional Activity Instructor Notes: ACTIVITY – SMALL GROUPS - Conduct a Module 2 culminating group activity with participants called “Top Ten.” - Divide the class into groups of 3-5, depending on the number of participants in the class. - Hand out a sheet of the flipchart paper to each group. - Ask participants to individually take 5 minutes to write down the top ten things they learned during this module. - Ask participants to write each item on a single post-it note, then arrange notes in order from most important to less important. - Once done, ask participants to review their post-it notes with the other group members for 10 minutes, then take 5 minutes to write out a group top ten list on their flipchart paper. - Tell participants that as a method for identifying and sorting items, they may wish to group similar post-it notes from each person to identify those items in common, then proceed to separating out items that are different. - Encourage groups to post their group top ten lists on the flipchart paper at the front of the class. - Proceed to the Activity Debrief on the next slide. |
| --- | --- |

| Slide 53 | Activity Debrief Instructor Notes: - Ask each group for a volunteer to present their group’s top ten list to the rest of the class. - Pose the following questions to participants to discuss their learning experience during this exercise. ASK: - “Did you learn anything new? If so, what?” - “How might you apply what you learned today to your job or to your short or long-term goals?” - “How does this information promote value or benefit to your role on the job?” |
| --- | --- |

| Slide 54 | Module 2 Summary Instructor Notes: - To wrap-up Module 2, review the learning objectives. |
| --- | --- |