## Slide 1
Ambient Air Monitoring
(AMBM 102)
Notes: CLASSROOM SETUP:​
Check classroom layout​
Instructor name/date/course title on board or easel​
Binders, name cards, markers, agenda, handouts on desks​
(Title of handouts - CEM Student Reference Guide)​
Setup equipment and verify operation​
Setup display materials​
Locate bathrooms and emergency exits/procedures​
INTRODUCTORY REMARKS:​
Fill out name tents, tracking cards, sign-in sheet. "Please write legibly as you want your name to appear on your certificate. Include company name and address."​
Agenda highlights - lecture, breaks, lunch, exam, field inspection trip, safety gear requirements, certificate. Have students turn off cell phones and set beepers to stun.​
Introduce yourself, background, and experience.​
Student background, experience, expectations.​
Aims of 200-Series training program:​
• Inspector-oriented​
• Details of process/systems and practical inspection techniques​
• Facilitate communication and share experiences​
Topics to be covered - refer to agenda, handout, or overview slide ​
Student Course Manual - review contents, indicate how used ​
Handouts - indicate how will you be using them. Discuss show and tell materials.​
Additional information - instructor discretion​

## Slide 2: Course Overview
Regulations and Standards
Monitoring Networks
Station Siting
Instrumentation
Documentation
Data Handling
Quality Assurance
References and Resources

## Slide 3
Standards and Regulations

## Slide 4: EPA Responsibilities Under CAA
National Ambient Air Quality Standards (NAAQS)
  Identification
  Attainment
Toxic air pollutants
  Identification
  Control
Acid Rain
Pollution Index
PSD
Notes: The Ambient Air Monitoring Program.​
The Federal government first allotted money to states to research ambient air was as a result of the Air Pollution Control Act of 1955 in section 103. Many states were already monitoring. The role of the federal government in coordinating this monitoring came as a result of the Air Pollution Control Act amendments of 1963. At that time the federal government through the Public Health Service began requiring specific types of equipment.​
​
In 1970, ammendments and now called the Clean Air Act (CAA) was signed into law. The CAA and its amendments provides the framework for all pertinent organizations to protect air quality. EPA's principal responsibilities under the CAA, include:​
Setting National Air Quality Standards (NAAQS) for pollutants considered harmful to the public health and environment , These six pollutants, also called criteria pollutants, are: particulate matter, sulfur dioxide, carbon monoxide, nitrogen dioxide, ozone, and lead.​
​
Ensuring the air quality standards are met or attained (in cooperation with the States) through national standards and strategies to control air emission standards from sources ​
Ensuring the sources of toxic air pollutants are well controlled ​
Monitoring the effectiveness of the program.​
​
​
​

## Slide 5: Monitoring
Attainment
Progress Toward Attainment
Pollution Trends
Emergency Control Procedures
Database
Notes: Monitor to determine attainment of NAAQS and California standards. California standards are often more stringent, more areas would be non-attainment. The differences in the standards will be discussed later. (REF: Student Handbook, pages 15-20)​
If non-attainment, need to determine progress toward attainment.​
Types of standards:​
NAAQS: National Ambient Air Quality Standards Primary NAAQS: protect public health​
Secondary NAAQS: protect public welfare​
California air quality standards: CA version of NAAQS, generally stricter​

## Slide 6: Regulations
U.S. EPA
  40 CFR 50 - NAAQS
  40 CFR 53 - Methods
  40 CFR 58 - Monitoring criteria
  40 CFR 52.21 – PSD
State and Local Regs
Notes: USEPA regulations:​
40 CFR 50: National Ambient Air Quality Standards (NAAQS).​
40 CFR 53: Ambient Air Monitoring Reference and Equivalent Methods.​
40 CFR 58: Monitoring criteria and station configurations.​
40 CFR 52.21: Prevention of Significant Deterioration (PSD). Student’s copy may say 51.24. That’s wrong. It was never that.​
California regulations (H&S Code)
Test Question #1

## Slide 7: Monitoring Networks
SLAM: State and Local Air Monitoring Station
NAMS: National Air Monitoring Station
PAMS: Photochemical Assessment Monitoring Station
Ncore: National Core Multipollutant Network
Notes: The EPA's ambient air quality monitoring program is carried out by State and local agencies and consists of the above major categories of monitoring stations. Most are explained on the following slides.​
PSD: No extra slide -- “PSD station means any station operated for the purpose of establishing the effect on air quality of the emissions from a proposed source for purposes of Prevention of Significant Deterioration as required by subpart 51.24(n) of part 51 of this chapter.” (Additional information: 40 CFR 58.1(d)).​
SPM: no extra slide -- Special Purpose Monitoring Stations provide for special studies needed by the State and local agencies to support State implementation plans and other air program activities. The SPMS are not permanently established and, can be adjusted easily to accommodate changing needs and priorities. The SPMS are used to supplement the fixed monitoring network as circumstances require and resources permit. If the data from SPMS are used for SIP purposes, they must meet all QA and methodology requirements for SLAM monitoring
Test Questions
#12-16

## Slide 8: Monitoring Networks
PSD: Prevention of Significant Deterioration
SPM: Special Purpose Monitoring
IMPROVE: Interagency Monitoring of Protected Visual Environments
Near Road NO2
Notes: The EPA's ambient air quality monitoring program is carried out by State and local agencies and consists of the above major categories of monitoring stations. Most are explained on the following slides.​
PSD: No extra slide -- “PSD station means any station operated for the purpose of establishing the effect on air quality of the emissions from a proposed source for purposes of Prevention of Significant Deterioration as required by subpart 51.24(n) of part 51 of this chapter.” (Additional information: 40 CFR 58.1(d)).​
SPM: no extra slide -- Special Purpose Monitoring Stations provide for special studies needed by the State and local agencies to support State implementation plans and other air program activities. The SPMS are not permanently established and, can be adjusted easily to accommodate changing needs and priorities. The SPMS are used to supplement the fixed monitoring network as circumstances require and resources permit. If the data from SPMS are used for SIP purposes, they must meet all QA and methodology requirements for SLAMS monitoring.​
Near Road NO2 – Started under the 2010 NO2 NAAQS review. Most of the sites are now multipollutant sites with PM, black carbon, CO or other contaminant monitors co-located.
Test Questions
#12-16

## Slide 9: State and Local Monitoring (SLAMS) Network
Notes: The SLAMS consist of a network of ~ 4,000 monitoring stations whose size and distribution is largely determined by the needs of State and local air pollution control agencies to meet their respective State implementation plan (SIP) requirements.​
“The SLAMS make up the ambient air monitoring network which is required by subpart 58.20 to be provided for in the state’s implementation plan. This definition places no restrictions on the use of the physical structure or facility housing the SLAMS. Any combination of SLAMS and any other monitors (Special Purpose, NAMS, PSD) may occupy the same facility or structure without affecting the respective definitions of those monitoring stations.” ​
s. These six primary SLAMS objectives are​
as follows: ​
(1) to determine highest concentrations expected to occur in the area covered by the​
network; ​
(2) to determine representative concentrations in areas of high population density;​
(3) to determine the impact on ambient pollution levels of significant sources or source​
categories; ​
(4) to determine general background concentration levels;​
(5) to determine the extent of Regional pollutant transport among populated areas, and​
in support of secondary standards; and​
(6) to determine the welfare-related impacts in more rural and remote areas (such as​
visibility impairment and effects on vegetation).​
The goal in designing the SLAMS networks is to establish monitoring stations that will provide data​
to meet these monitoring objectives and as input to EPA and States to improve environmental​
decisions.

## Slide 10: National Air Monitoring (NAMS) Network
Notes: The NAMS (1,080 stations) are a subset of the SLAMS network with emphasis being given to urban and multi-source areas.​
In effect, they are key sites under SLAMS, with emphasis on areas of maximum concentrations and high population density.​
“Collectively the NAMS are a subset of the SLAMS ambient air quality monitoring network.” (Additional information: 40 CFR 58.1(c), direct quote, emphasis added) NAMS are usually in areas of special interest.​
The primary objective of the NAMS is to monitor in areas where the​
pollutant concentration and population exposure are expected to be the highest consistent with the​
averaging time of the NAAQS. The NAMS are a subset of SLAMS that focus on urban and​
multisource areas

## Slide 11: Photochemical Assessment Monitoring (PAMS) Network
Updated 08/24
Notes: A PAMS network is required for each O3 nonattainment area that is designated serious, severe, or extreme. Ozone and its precursors (VOC and Nox) plus met data are to be monitored.​
As mandated by the 1990 Clean Air Act Amendments, EPA is requiring more extensive monitoring of ozone and its precursors in areas with persistently high ozone levels (mostly large cities). The States must establish air monitoring sites, called Photochemical Assessment Monitoring Stations, or PAMS, to collect and report detailed data for volatile organic compounds, nitrogen oxides, ozone, and meteorology. Analysis of the data will help EPA and the States to better understand the underlying causes of ozone pollution, devise effective remedies, and measure improvement. ​
Many of the methods and procedures are new, requiring quality assurance assistance to the State and local agencies by EPA. It is important that coordination be maintained between the regional offices and with Headquarters in order to provide consistency between PAMS sites. Several Regions have agreed to have members serve on a national PAMS Quality Assurance Work Group assembled to discuss QA/QC procedures which can be uniformly and inexpensively applied to PAMS. ​
The required networks will have from two to five sites, depending on the population of the area. There will be a phase-in period of one site per year starting in 1994. The ultimate PAMS network could exceed 90 sites at the end of the 5-year phase-in period.

## Slide 13: NCore Network
Updated 08/24

## Slide 14: NCore Objectives
Timely reporting of air quality data to public
Support for development of emission strategies
Long-term tracking of emission strategies
Long-term health assessment for NAAQS reviews
Establish Attainment/Nonattainment areas
Support for scientific studies in technical, health & atmospheric disciplines
Support to ecosystem assessment
Notes: ​Timely reporting of data to public by supporting AIRNow, air quality forecasting, and other public reporting mechanisms;​
Support for development of emission strategies through air quality model evaluation and other observational methods;​
Accountability of emission strategy progress through tracking long-term trends of criteria and non-criteria pollutants and their precursors;​
Support for long-term health assessments that contribute to ongoing reviews of the NAAQS;​
Compliance through establishing nonattainment/attainment areas through comparison with the NAAQS;​
Support to scientific studies ranging across technological, health, and atmospheric process disciplines; and​
Support to ecosystem assessments recognizing that national air quality networks benefit ecosystem assessments and, in turn, benefit from data specifically designed to address ecosystem analyses.

## Slide 15: NCore pollutants

## Slide 16
IMPROVE Sites
Notes: The Interagency Monitoring of Protected Visual Environments (IMPROVE) steering committee is composed of representatives from EPA, the National Park Service(NPS), USDA Forest Service (USFS), Fish and Wildlife Service (FWS), the Bureau of Land Management (BLM), State and Territorial Air Pollution Program Administrators (STAPPA), Western States Air Resources Cooperative (WESTAR), and Northeast States for Coordinated Air Use Management (NESCAUM).​
A network has been established to collect visibility related data associated primarily with Class I areas. The IMPROVE Newsletter may be found in the News Section of AMTIC.

## Slide 17
Texas Commission on Environmental Quality 2023 Ambient Monitoring Network

## Slide 18
NATTS Monitors

## Slide 19
Criteria Pollutants

## Slide 20: Current NAAQS
9.0 µg/m3

## Slide 21: Carbon Monoxide(CO) Standards-Table of Historical CO NAAQS
Final Rule/Decision | Primary/ Secondary | Indicator | Averaging Time | Level (2) | Form
1971 36 FR 8186 Apr 30, 1971 | Primary and Secondary | CO (1) | 1-hour period | 35 ppm | Maximum, not to be exceeded more than once in a year
 |  |  | 8-hour period | 9 ppm | Maximum, not to be exceeded more than once in a year (3)
1985 50 FR 37484 Sept 13, 1985 | Primary standards retained, without revision; secondary standards revoked. |  |  |  | 
1994 59 FR 38906 Aug 1, 1994 | Primary standards retained, without revision. |  |  |  | 
2011 76 FR 54294 Aug 31, 2011 | Primary standards retained, without revision. |  |  |  | 

## Slide 22
CO
Current CO Monitoring
Notes: Note: On this and the following AIRS slides the shaded states have monitors.​
The ARB maintains 23 CO sites with 8 in Mexico and the districts maintain 68. ​
Health and Environmental Effects: Carbon monoxide enters the bloodstream and reduces oxygen delivery to the body's organs and tissues. The health threat from exposure to CO is most serious for those who suffer from cardiovascular disease. Healthy individuals are also affected, but only at higher levels of exposure. ​
Exposure to elevated CO levels is associated with visual impairment, reduced work capacity, reduced manual dexterity, poor learning ability, and difficulty in performing complex tasks. ​

## Slide 23: Nitrogen Dioxide (NO2) Standards-Table of Historical NO2 NAAQS
Final Rule/Decision | Primary/ Secondary | Indicator (1) | Averaging Time | Level (2) | Form
1971 36 FR 8186 Apr 30, 1971 | Primary and Secondary | NO2 | Annual | 53 ppb (3) | Annual arithmetic average
1985 50 FR 25532 Jun 19, 1985 | Primary and secondary NO2 standards retained, without revision. |  |  |  | 
1996 61 FR 52852 Oct 8, 1996 | Primary and secondary NO2 standards retained, without revision. |  |  |  | 
2010 75 FR 6474 Feb 9, 2010 (4) | Primary | NO2 | 1-hour | 100 ppb | 98th percentile, averaged over 3 years (5)
 |  | Primary annual NO2 standard retained, without revision. |  |  | 
2018 83 FR 17226 April 8, 2018 | Primary | Primary annual NO2 standard retained, without revision. |  |  | 
Notes: ​

## Slide 24
Current NO2 Monitors
Updated 08/24

## Slide 25: Oxides of Sulfur (SO2) Standards-Table of Historical SO2 NAAQS
Final Rule/Decision | Primary/ Secondary | Indicator (1) | Averaging Time | Level (2) | Form
1971 36 FR 8186 Apr 30, 1971 | Primary | SO2 | 24-Hour | 0.14 ppm | Not to be exceeded more than once per year
 |  |  | Annual | 0.03 ppm | Annual arithmetic average
 | Secondary |  | 3-Hour | 0.5 ppm | Not to be exceeded more than once per year
 |  |  | Annual (3) | 0.02 ppm | Annual arithmetic average
1973 38 FR 25678 Sept 14, 1973 | Secondary | Secondary 3-hour SO2 standard retained, without revision; secondary annual SO2 standard revoked. |  |  | 
1996 61 FR 25566 May 22, 1996 | Primary | Existing primary SO2 standards retained, without revision. |  |  | 
2010 75 FR 35520 Jun 22, 2010 (4) | Primary | SO2 | 1-hour | 75 ppb | 99th percentile, averaged over 3 years (5)
 |  | Primary annual and 24-hour SO2 standards revoked. |  |  | 
77 FR 20218 April 3, 2012 | Secondary | Existing secondary SO2 standard (3-hour average) retained, without revision. |  |  | 
84 FR 9866 March 18, 2019 | Primary | Existing primary SO2 standard retained, without revision. |  |  | 

## Slide 26
Current SO2 Monitors
Notes: The ARB maintains 11 SO2 sites including 8 in Mexico while the districts maintain 48. ​
Health and Environmental Effects: The major health concerns associated with exposure to high concentrations of SO2 include effects on breathing, respiratory illness, alterations in pulmonary defenses, and aggravation of existing cardiovascular disease. Children, the elderly, and people with asthma, cardiovascular disease or chronic lung disease (such as bronchitis or emphysema), are most susceptible to adverse health effects associated with exposure to SO2 .​
SO2 is a precursor to sulfates, which are associated with acidification of lakes and streams, accelerated corrosion of buildings and monuments, reduced visibility, and adverse health effects.​

## Slide 27
SO2 Nonattainment Areas

## Slide 28
SO2 Nonattainment Areas (1971 Std)
Notes: Trends in SO2 Levels: Between 1986 and 1995, national SO2 concentrations decreased 37 percent and SO2 emissions decreased 18 percent. Between 1994 and 1995, national SO2 concentrations decreased 17 percent and SO2 emissions decreased 13 percent. These significant decreases in concentrations and emissions reflect the success of the first year of the Acid Rain Program. ​
While national SO2 air quality levels have improved, EPA remains concerned about short-term peak SO2 concentrations. As a result, EPA proposed a program for States, industry, and communities to use in evaluating and addressing peak concentrations that could occur near some industrial sources.

## Slide 29: Lead (Pb) Standards- Table of Historical Pb NAAQS
Final Rule/Decision | Primary/ Secondary | Indicator | Averaging Time | Level (1) | Form
1978 43 FR 46246 Oct 5, 1978 | Primary and Secondary | Pb-TSP (2) | Calendar Quarter | 1.5 µg/m3 | Not to be exceeded
Feb 21, 1991 – Agency released multimedia “Strategy for Reducing Lead Exposures” (3) |  |  |  |  | 
2008 73 FR 66964 Nov 12, 2008 | Primary and Secondary | Pb-TSP | 3-month period | 0.15 µg/m3 | Not to be exceeded
2016 81 FR 71906 Oct 18, 2016 | Primary and secondary standards retained, without revision. |  |  |  | 
Notes: Along with a change in the NAAQS the monitoring networks were changed requirng Pb monitoring at an expanded number of source sites and at larger municipal airports​
Merrill Field ................................................ Anchorage ..AK ​
Pryor Field Regional .................................... Limestone ..AL ​
Palo Alto Airport of Santa Clara County ......... Santa Clara .CA ​
McClellan-Palomar ...................................... San Diego …CA​
Reid-Hillview ..............................................Santa Clara ..CA ​
Gillespie Field .............................................San Diego …CA ​
San Carlos .................................................. San Mateo ..CA ​
Nantucket Memorial .................................... Nantucket …MA ​
Oakland County International ....................... Oakland .....MI ​
Republic ......................................................Suffolk .......NY ​
Brookhaven ................................................ Suffolk .......NY ​
Stinson Municipal ........................................ Bexar ........ TX ​
Northwest Regional ..................................... Denton ......TX ​
Harvey Field ................................................ Snohomish ..WA ​
Auburn Municipal ......................................... King ..........WA

## Slide 30
Pb
Current Lead Monitors
Updated 08/24
Notes: The ARB maintains 15 lead sites (including 9 in Mexico), 1 collocated site in Bakersfield, and the districts maintain 18.​
Health Effects: Exposure to lead mainly occurs through inhalation of air and ingestion of lead in food, paint, water, soil, or dust. Lead accumulates in the body in blood, bone, and soft tissue. Because it is not readily excreted, lead can also affect the kidneys, liver, nervous system, and other organs. Excessive exposure to lead may cause anemia, kidney disease, reproductive disorders, and neurological impairments such as seizures, mental retardation, and/or behavioral disorders. ​
Even at low doses, lead exposure is associated with changes in fundamental enzymatic, energy transfer, and other processes in the body. Fetuses and children are especially susceptible to low doses of lead, often suffering central nervous system damage or slowed growth. Recent studies show that lead may be a factor in high blood pressure and subsequent heart disease in middle-aged white males. Lead may also contribute to osteoporosis in post-menopausal women.

## Slide 31
Lead Nonattainment Areas

## Slide 32
Notes: Trends in Lead Levels: Between 1986 and 1995, average lead concentrations in urban areas throughout the country decreased 78 percent while total lead emissions decreased 32 percent. These reductions are a direct result of the use of unleaded gasoline in automobiles. The large reduction in lead emissions from transportation sources has changed the nature of the pollution problem for lead in the U.S. ​
While there are still violations of the lead air quality standard, they tend to occur near large industrial sources such as lead smelters. Between 1994 and 1995, lead emissions decreased 1 percent while national average lead concentrations remained unchanged.

## Slide 33
Pb
Notes: Control techniques include:​
Reducing or eliminating lead in processes and products such as fuels and coatings.​

## Slide 34: Ozone (O3) Standards-Table of Historic O3 NAAQS
Final Rule/Decision | Primary/ Secondary |  | Indicator (1) | Averaging Time | Level (2) | Form
1971 36 FR 8186 Apr 30, 1971 | Primary and Secondary |  | Total photochemical oxidants | 1-hour | 0.08 ppm | Not to be exceeded more than one hour per year
1979 44 FR 8202 Feb 8, 1979 | Primary and Secondary |  | O3 | 1-hour | 0.12 ppm | Attainment is defined when the expected number of days per calendar year, with maximum hourly average concentration greater than 0.12 ppm, is equal to or less than 1
1993 58 FR 13008 Mar 9, 1993 | EPA decided that revisions to the standards were not warranted at the time |  |  |  |  | 
1997 62 FR 38856 Jul 18, 1997 | Primary and Secondary | O3 | O3 | 8-hour | 0.08 ppm | Annual fourth-highest daily maximum 8-hr concentration, averaged over 3 years
2008 73 FR 16483 Mar 27, 2008 | Primary and Secondary | O3 | O3 | 8-hour | 0.075 ppm | Annual fourth-highest daily maximum 8-hr concentration, averaged over 3 years

## Slide 36
OZONE Formation
Notes: Ozone formation patterns from 8:00 am to 11:00 pm.

## Slide 37: Ground-Level Ozone
Notes: Sacramento Ground Level Ozone​
from 8am to 7pm

## Slide 39
Current Ozone Monitors
Updated 08/24
Notes: The ARB maintains 37 O3 sites with 8 in Mexico while the districts maintain 151.​
Health and Environmental Effects: Scientific evidence indicates that ground-level ozone not only affects people with impaired respiratory systems (such as asthmatics), but healthy adults and children as well. Exposure to ozone for 6 to 7 hours, even at relatively low concentrations, significantly reduces lung function and induces respiratory inflammation in normal, healthy people during periods of moderate exercise. It can be accompanied by symptoms such as chest pain, coughing, nausea, and pulmonary congestion. ​
Recent studies provide evidence of an association between elevated ozone levels and increases in hospital admissions for respiratory problems in several U.S. cities. Results from animal studies indicate that repeated exposure to high levels of ozone for several months or more can produce permanent structural damage in the lungs. ​
Ozone is responsible for approximately 1 to 2 billion dollars of agricultural crop yield loss in the U.S. each year. Ozone also damages forest ecosystems in California and the eastern U.S. New scientific studies indicate that ozone causes adverse health and environmental effects at lower concentrations and longer periods of exposure than the current standards. ​

## Slide 40
Ozone Nonattainment Areas (2008 Std)
(8 hour, average 4th maximum, 0.75 ppm)
Notes: The old one hour standard applies to present non-attainment areas. It no longer applies to an area once the EPA determines that an area has air quality standards that meet the one-hour standard. At that point the new 8 hour standard applies.​
(See 40 FR Part 50 -- 50.9 and 50.10)​
Control strategies:​
Reduce Nox (already discussed)​
Reduce VOC:​
Solvents​
Paints & other coatings​
Vapor recovery​
Control devices - carbon beds, incineration, bioremediation

## Slide 41
O3

## Slide 42: Particulate Matter (PM) Standards - Table of Historical PM NAAQS
Final Rule | Primary/ Secondary | Indicator | Averaging Time | Level (1) | Form
1971 36 FR 8186 Apr 30, 1971 | Primary | TSP (2) | 24-hour | 260 µg/m3 | Not to be exceeded more than once per year
 |  |  | Annual | 75 µg/m3 | Annual Average
 | Secondary | TSP | 24-hour | 150 µg/m3 | Not to be exceeded more than once per year
1987 52 FR 24634 Jul 1, 1987 | Primary and Secondary | PM10 | 24-hour | 150 µg/m3 | Not to be exceeded more than once per year on average over a 3-year period
 |  |  | Annual | 50 µg/m3 | Annual arithmetic mean, averaged over 3 years
1997 62 FR 38652 Jul 18, 1997 | Primary and Secondary | PM2.5 | 24-hour | 65 µg/m3 | 98th percentile, averaged over 3 years
 |  |  | Annual | 15.0 µg/m3 | Annual arithmetic mean, averaged over 3 years (3),(4)
 |  | PM10 | 24-hour | 150 µg/m3 | Initially promulgated 99th percentile, averaged over 3 years; when 1997 standards for PM10 were vacated, the form of 1987 standards remained in place (not to be exceeded more than once per year on average over a 3-year period) (5)
 |  |  | Annual | 50 µg/m3 | Annual arithmetic mean, averaged over 3 years
2006 71 FR 61144 Oct 17, 2006 | Primary and Secondary | PM2.5 | 24-hour | 35 µg/m3 | 98th percentile, averaged over 3 years (6)
 |  |  | Annual | 15.0 µg/m3 | Annual arithmetic mean, averaged over 3 years (2), (7)
 |  | PM10 | 24-hour | 150 µg/m3 | Not to be exceeded more than once per year on average over a 3-year period
2012 | Primary | PM2.5 | Annual | 12.0 µg/m3 | Annual arithmetic mean, averaged over 3 years
 |  |  | Annual | 15.0 µg/m3 | Annual arithmetic mean, averaged over 3 years
 | Secondary |  |  |  | 
 | Primary and Secondary |  | 24-hour | 35 µg/m3 | 98th percentile, averaged over 3 years
 | Primary and Secondary | PM10 | 24-hour | 150 µg/m3 | Not to be exceeded more than once per year on average over a 3-year period

## Slide 43: PM Standards Revision in 2024
Final Rule | Primary/ Secondary | Indicator | Averaging Time | Level (1) | Form
89 FR 16202 | Primary | PM2.5 | Annual | 9.0 µg/m3 | Annual arithmetic mean, averaged over 3 years 2, 7
2024 | Secondary PM2.5 standards, and primary and secondary PM10 standards, retained without revision. |  |  |  | 

## Slide 44
Current PM 10 and 2.5 Monitors
Updated 08/24

## Slide 46
PM-10 Nonattainment Areas

## Slide 47
PM-2.5 Nonattainment Areas

## Slide 48
PM-2.5 Nonattainment Areas
For Standards of Different Years

## Slide 51
Siting Criteria

## Slide 52: Network Design Considerations
Concentration Expected
Representative Concentrations
Significant Sources or Source Categories
Notes: Monitoring objectives: Why are we monitoring? What are we looking for? For example: measure a source impact (also need a background concentration monitor for comparison).​
These objectives indicate the nature of the samples that the monitoring network will collect and will be used during the development of data quality objectives (section 3). As one reviews the objectives it becomes apparent that it will be rare that sites can be located to meet more than two or three objectives. Therefore, organization need to prioritize their objectives in order to choose the sites that most representative of that objective and therefore will provide data of adequate quality.​
Through the process of implementing the CAA, a number of ambient air quality monitoring networks have been developed. The EPA's Ambient Air Quality Monitoring Program is carried out by State and local agencies and consists of four major categories of monitoring stations or networks that measure the criteria pollutants.
10:00am BREAK

## Slide 53: Network Design Considerations
Background Concentrations
Regional Transport
Welfare-Related Impacts for Rural Areas
Notes: Monitoring objectives: Why are we monitoring? What are we looking for? For example: measure a source impact (also need a background concentration monitor for comparison).​
These objectives indicate the nature of the samples that the monitoring network will collect and will be used during the development of data quality objectives (section 3). As one reviews the objectives it becomes apparent that it will be rare that sites can be located to meet more than two or three objectives. Therefore, organization need to prioritize their objectives in order to choose the sites that most representative of that objective and therefore will provide data of adequate quality.​
Through the process of implementing the CAA, a number of ambient air quality monitoring networks have been developed. The EPA's Ambient Air Quality Monitoring Program is carried out by State and local agencies and consists of four major categories of monitoring stations or networks that measure the criteria pollutants.
10:00am BREAK

## Slide 54: Scales of Monitoring
Micro
Middle
Neighborhood
Notes: Micro​
Concentrations in Air Volumes from Several Meters to about 100 Meters​
Usually High Concentration Impact​
Middle​
100 Meters to about 0.5 Kilometers​
Often Source Impact​
Neighborhood​
0.5 to 4 Kilometers​
Often Population Impact.

## Slide 55: Micro Scale Site Up to 100 m
Notes: (TQ #19)​
​
Example map of micro scale site (South Lake Tahoe)​
​
A microscale site is to measure localized pollutant concentrations with essentially no mixing (for example, high CO concentrations beside a street). ​
​
There is a microscale CO monitoring station at this location. The inlet hangs over the sidewalk beside the street.

## Slide 56: Middle Scale Site 100 m to 0.5 km
Notes: Example map of middle scale site (South Lake Tahoe)​
Represents a somewhat larger area; several blocks around.​
Used to measure high concentrations and source impacts with minimal mixing.​
​
​
((SOURCE IMPACT))

## Slide 57: Neighborhood Scale Site 0.5 km to 4 km
Notes: Example map of neighborhood scale site (South Lake Tahoe)​
Neighborhood scale is the most common monitoring scale. Ozone impact areas are often neighborhood scale.​
Used to measure population and source impacts with moderate mixing.

## Slide 58: Additional Scales of Monitoring
Urban
Regional
National and Global
Notes: Urban​
4 to 50 Kilometers​
Usually Population Impacts​
Regional​
Tens to Hundreds of Kilometers​
Often Urban Transport Impact Area​
National and Global​
Concentrations in Air Volumes Representing a Nation or World as a Whole​
Usually Background Concentrations.

## Slide 59: Urban Scale Site 4 km to 50 km
Notes: Example map of urban scale site (South Lake Tahoe)​
Urban scale would be about the size of the lake in the Lake Tahoe basin and would represent pollutants that are well mixed throughout the basin. ​
Used to monitor pollutants distributed over large areas, often pollutants transported from urban areas. ​
Ozone impact areas downwind of cities are often urban scale.​
Useful for assessing air quality over an entire metropolitan area.​
Particulate sulfates are well mixed and uniform throughout the Tahoe basin in this example slide.

## Slide 60: Regional Scale Site 10s km to 100s km
Notes: Example map of regional scale sites​
Used to monitor well mixed, often transported, pollutants over large areas and for background concentrations.​
​
​
URBAN TRANSPORT

## Slide 61: National & Global Scale Site
100s to
1,000s km
(On the Scale
of a Nation
or the World as a Whole)
Notes: ​
Example map of national or global scale site​
An example would be Mona Loa in Hawaii. Mona Loa can measure particulate from the Gobi Desert and soot from coal burning in Beijing. It is also the station used to monitor global CO2 levels (global warming greenhouse effect predictions are based on these results).​
​
​

## Slide 62: Monitoring Objectives & Scale
Monitoring Objective
Appropriate Siting Scales
Highest concentration Source impact
Population
General / Background
Micro, middle, neighborhood (sometimes urban)
Micro, middle, neighborhood
Neighborhood, urban
Neighborhood, regional, global
Notes: For specific objectives stations need corresponding scales.​
High concentrations - fairly small (localized) scale to find the ‘hot spot’.​
Background levels - monitor large, well mixed, areas.​

## Slide 64: Network Design Considerations
Priority area (zone of highest pollution conc.)
Air Transport
Evaluation
Notes: The information required for selecting the number of samplers and the sampler locations include isopleth maps, population density maps, and source locations. The following are suggested guidelines:​
​
1) The priority area is the zone of highest pollution concentration within the region; one or more stations are to be located in this area.​
​
2) The quality of air entering the region is to be assessed by stations situated on the periphery of the region; meteorological factors (e.g., frequencies of wind directions) are of primary importance in locating these stations.​
​
3) A major objective of surveillance is evaluation of progress made in attaining the desired air quality; for this purpose, sampling stations should be strategically situated to facilitate evaluation of the implemented control tactics

## Slide 65: Network Design Considerations
Population Areas
Future development
Full Representation
Notes: Close attention should be given to densely populated areas​
​
Sampling should be undertaken in areas of projected growth to determine the effects of future development​
​
All portions of the region should be represented​
​
Some stations will be capable of fulfilling more than one of the functions indicated; for example, a station located in a densely populated area can indicate population exposures and can also document the changes in pollutant concentrations resulting from mitigation strategies used in the area.

## Slide 66: Number of Stations – PM10
1 Exceeding NAAQS by 20% or more, or 95% Probability of PM10 Nonattainment
2 Exceeding 80% of NAAQS, or 20% to 95% Probability of PM10 Nonattainment
3 Less than 80% NAAQS, or < 20% Probability of PM10 Nonattainment
Population
Expected Maximum Concentration
High1
Medium2
Low3
> 1,000,000
500,000 - 1,000,000
250,000 - 500,000
100,000 - 250,000
6 - 10 4 - 8 2 - 4
4 - 8 2 - 4 1 - 2
3 - 4 1 - 2 0 - 1
1 - 2 0 - 1 0
Notes: (Additional information: 40 CFR 58, Appendix D)​
How many monitoring stations are needed? The EPA says this many NAMS for PM10 should be setup .​
Similar tables are in the CFR for other pollutants.​
(Additional information: 40 CFR 58)​
PAMS stations are a little different: PAMS stations are arrayed around urban areas in specified patterns (tables in CFR). (Additional information: 40 CFR 58)​
​
These last two charts are just examples. Additional ones (including PM 2.5) are in the CFR.​

## Slide 67: Station Siting Considerations
Available sites
Start-up costs
  Equipment
  Facility improvements
Operation costs
  Equipment operation and maintenance
  Station costs (lease payments, heating, etc.)
  Expendables (calibration gases, chart paper, etc.)
  Personnel
Notes: ​
After deciding what pollutants to monitor and where to monitor. Find a place for the station.​
Many stations are in trailers that can be brought into a site. Many also are in lease space in strip shopping centers. Sometimes space in the district offices, state offices, city offices, or county offices can be arranged. Schools and fire houses are often good sites because they are seldom located on main roads.​
Costs of a monitoring station can also be substantial (REF: Personal communication from MLD monitoring section managers)​
Equipment costs: monitors are $5,000-20,000 each; a station can easily cost over $100,000 to setup.​
Lease space: in many areas the lease space can be quite expensive, and is a continuing expense.​
Technician: can maintain one main station and several satellite stations.​
Calibration gases, strip chart paper, and other expendables​
Equipment maintenance

## Slide 68: Station Siting Considerations
Types of Pollutants
Topography
Air flow
Notes: Station siting is very important - If a station is not properly located the data generated is useless (or worse, misleading). (REF: Student Handbook, page 96-137)​
​
​
Garbage In - Garbage Out​

## Slide 69: Station Categories
A (Ground Level) | Heavy pollutant concentration, high potential for pollutant buildup
B (Ground Level) | Heavy pollutant concentration, minimal potential for buildup
C (Ground Level) | Moderate pollution concentration
D (Ground Level) | Low pollutant concentration
E (Air Mass) | Sampler probe that is between 6-45m (20-150 ft) above ground
F (Source-Oriented) | Sampler that is adjacent to a point source
Notes: Station Category /Characterization​
A (ground level) Heavy pollutant concentrations, high potential for pollutant buildup. A site 3 to 5 m (10-16 ft) from major traffic artery and that has local terrain features restricting ventilation. A sampler probe that is 3 to 6 m (10-20 ft) above ground.​
B (ground level) Heavy pollutant concentrations, minimal potential for a pollutant buildup. A site 3 to 15 m (15-50 ft) from a major traffic artery, with good natural ventilation. A sampler probe that is 3 to 6 m (10-20 ft) above ground.​
C (ground level) Moderate pollutant concentrations. A site 15 to 60 m (5-200 ft) from a major traffic artery. A sampler probe that is 3 to 6 m ( 10-20 ft ) above ground.​
D (ground level) Low pollutant concentrations. A site 60 > m (> 200 ft) for a traffic artery. A sampler probe that is 3 to 6 m (10-20 ft) above ground.​
E (air mass) Sampler probe that is between 6 and 45 m (20-150 ft) above ground. Two subclasses: (1) good exposure from all sides (e.g., on top of building) or (2) directionally biased exposure (probe extended from window).​
F (source-oriented) A sampler that is adjacent to a point source. Monitoring that yields data directly relatable to the emission source.​

## Slide 70: Site Information
Local Sources
  Flues & Vents by Inlet
  Non-Vehicular/Local Industry
  Traffic
Dominant Influence Category
  Industrial
  Residential
  Commercial
  Vehicular
  Urbanization
  Near Urban
  Agricultural
  Recreational Area
Notes: Appendix E of Part 58 has all the specifics but for an idea ​
​
​
Check for local sources of pollutants that will interfere with the objective of the station.​
Vents and flues on the roof near the inlet (check for furnace and water heater vents, emergency generator stacks, etc.).​
Local industry including filling stations, dry cleaners, auto shops, etc.​
Traffic check distance to streets and traffic volume. Check traffic table for each pollutant.​
Dominant Influence Category summarizes the sources. There are seven categorize (plus an ‘other’ category), choose one: industrial, residential, commercial, vehicular, near urban, agricultural, or recreational area.​
​

## Slide 71
Local Sources Near Monitoring Stations
Notes: The success of the PAMS monitoring program is predicated on the fact that no site is unduly influenced by any one stationary emissions source or small group of emissions sources. Any significant influences would cause the ambient levels measured by that particular site to mimic the emissions rates of this source or sources rather than following the changes in nonattainment area-wide emissions as intended by the Rule.​
For purposes of this screening procedure, if more than 10% of the typical "lower end" concentration measured in an urban area is due to a nearby source of precursor emissions, then the PAMS site must be relocated or a more refined analysis conducted than is presented here. Detailed procedures can be fund in the EPA PAMS Implementation Manual . ​
​

## Slide 72: Site Information
Data Acquisition Objective
Station Type
Spatial Scale
Instrumentation
Sampling System
Influential Pollutant Sources
Topography
Atmospheric Exposure
Notes: A sampling site or an array of sites for one pollutant may be appropriate for another pollutant species because of the configuration of sources, the local meteorology, or the terrain. Pollutants undergo changes in their compositions between their emission and their detection; therefore, the impact of that change on the measuring system should be considered. ​
Atmospheric chemical reactions such as the production of O3 in the presence of NO and Hydrocarbons (HC), and the time delay between the emission of NOx and HC and the detection peak of O3 values may require either a sampling network for the precursors of O3 and/or a different network for the actual O3 measurement.

## Slide 73: Site Information
Obstacles
  Description
  Distance
  Height above inlet
  Walls
  Air flow arc
Trees
  As obstacles
  As interferants
Notes: Describe any obstacles and trees around the inlet.​
Obstacles will prevent pollutants from getting to the inlet or disrupt even air flow. Must be twice the height from the inlet (30o angle) and clear air flow to the inlet.​
Trees can be obstacles and also act as interferants.​
May be source of resuspended particulate; source of hydrocarbons; sink for pollutants (particulate deposition surface); or reactive surface (remove ozone, etc.) ​
Other influential features include hills, valleys, and bodies of water.

## Slide 74: Obstacle Effects
Notes: Area of perturbed air flow around an obstacle can be quite extensive. ​
​
Airport landing is 95% with the prevailing winds

## Slide 75: Location of Monitors
C = Core site
S = SLAMS site
p = Special Purpose Monitor
Notes: Network for fine particulate matter​
Core network (C)​
population-oriented SLAMS​
background sites​
transport sites​
NAMS sites for long-term monitoring​
Core category ‘a’ sites - NAMS site in area of expected maximum concentration.​
Core category ‘b’ sites - NAMS located in area of poor air quality with maximum population impact.​
Other SLAMS monitors (S)​
Supplementary Special Purpose Monitors (p)​
Cover student handout on monitoring site problems

## Slide 76
Measurement Process

## Slide 77: Air Pollutant Measurement Process
Separate pollutant from air
Determine pollutant quantity and air volume
Calculate pollution concentration by dividing pollutant quantity by air volume
Analyze data
Notes: Ambient air sampling is primarily concerned with the atmospheric concentrations of such pollutants as particulates, SO2, NOx , CO, and photochemical oxidants. To establish the basic validity of such ambient air monitoring data, it must be shown that:​
The proposed sampling method complies with the appropriate testing regulations.​
The equipment is accurately sited.​
The equipment was accurately calibrated using correct and established calibration methods.​
The organization implementing the data collection operation are qualified and competent.​

## Slide 78: Types of Monitoring
Automated analytical methods
  Point analyzers
  Open path analyzers
Time averaged samplers
  Manual methods
  Filter (ex. PM10) samples
Notes: Automated methods - continuously samples and analyzes.​
Sampler - collects a sample which is then analyzed in a separate step (usually mailed to a central laboratory).​
A proper sampling environment demands control of all physical parameters external to the samples that might affect sample stability, chemical reactions within the sampler, or the function of sampler components.​
​

## Slide 79: Electromagnetic Spectrum
400 nm
700 nm
Visible Range
Notes: (T.Q. #6)
Wavelength in m or nm Frequency in Hertz​
​
Ultraviolet (UV) region to the left (<400 nm)​
 Short wavelength and high energy​
​
Visible light in the center (400 to 700 nm)​
​
Infrared (IR) region to the right (above 700 nm)​
 Longer wavelength and lower energy​
​
CEM Analyzers use from 200 nm (UV) - 6000 nm (IR)​
​
Light carries energy and in some reactions with matter​
 It behaves as if it were composed of discrete packets of energy called photons ​
​
Light of short wavelengths (UV) carry more (photons)​
Light of longer wavelengths (IR) ​
​
​

## Slide 80: Beer-Lambert Law
Absorption of light related to:
  Absorption coefficient dependencies
    Wavelength of light
    Properties of the pollutant molecule
  Number of molecules in light path
    Concentration
    Path length
Notes: What Mr Beer and Mr Lambert said that for each wavelength of light The light is absorbed going through the gas depending on how many molecules it encounters and over what length actually Mr Lambert only spoke of the length not the concentration that was Mr Beer ​
The value of the absorption coefficient is dependent upon :​
:​
​
  of light and ​
 Properties of the pollutant molecule ​
​
​
The coefficient expresses the degree to which :A molecule will absorb light energy at a given  for a given length ​
​
 ​

## Slide 82: Analytical Techniques
Infrared Methods
  Differential Absorption
  Gas Filter Correlation
  Fourier Transform Infrared
Ultraviolet Methods
  Differential Absorption
  Second Derivative Spectroscopy
Visible Light - Opacity Measurement
  Scattering & Absorption
Notes: Most monitors use some frequency of light:​
 in the analysis of pollutant gases​
3 interactions of light with atoms and molecules​
Differential absorption spectroscopy (DOAS) is used to measure concentrations of trace gases. When combined with basic optical spectrometers such as prisms or diffraction gratings and automated, ground-based observation platforms, what we have is a cheap and powerful means for the measurement of such trace gas species as ozone and nitrogen dioxide.​
Gas filter coorelation—CO—splits light through 0% CO then 100% then thru the sample​
Fourier Transform infrared-- a technique which is used to obtain an infrared spectrum ofabsorption, emission, photoconductivity or Raman scattering of a solid, liquid or gas. An FTIR spectrometer simultaneously collects spectral data in a wide spectral range. This confers a significant advantage over a dispersive spectrometer which measures intensity over a narrow range of wavelengths at a time.​
​
(T.Q. #7)

## Slide 83: Analytical Techniques
Luminescence Methods
  Fluorescence
  Chemiluminescence
  Flame Photometry
Electroanalytical Methods
  Polarography
  Electrocatalytic
  Paramagnetism
  Conductivity
Notes: Note: Although the EPA approves ambient air monitors​
Stack CEM systems are approved  Case-by-case​
Fluorescence spectroscopy aka fluorometry or spectrofluorometry, is a type of electromagnetic spectroscopy which analyzes fluorescence from a sample. It involves using a beam of light, usually ultraviolet light, that excites the electrons inmolecules of certain compounds and causes them to emit light; typically, but not necessarily, visible lig​
Chemiluminescence (sometimes "chemoluminescence") is the emission of light (luminescence), as the result of a chemical reaction. There may also be limited emission of heat. ​
◊ In theory, one photon of light should be given off for each molecule of reactant,.​
Polarography a subclass of voltammetry where the working electrode is a dropping mercury electrode (DME) or a static mercury drop electrode (SMDE), which are useful for their wide cathodic ranges and renewable surfaces.
T.Q. #7

## Slide 84: Site Information
Site Description
  Ground Cover
  Height of Inlet
  Type of Samplers
  Spacing Between Samplers
  Inlet Boom Description and Orientation
  Meteorological Instrument Tower Description
  Meteorological Instrument Radiation Shield
Notes: Purpose, objective, and scale:​
Previously discussed, record in spaces provided.​
Ground cover at the site: ​
Look for sources of contaminants (dusty ground around a PM10 sampler) or ground surfaces that would absorb or generate pollutants and interferants.​
Probe inlet height (or meteorological instrument location): ​
Height above ground important for scale.​
Spacing between inlets: ​
Prevents interference.​
Inlet boom length: ​
For inlets on the side of a building.​
Met tower: ​
Describe the tower and placement - can interfere with some meteorological instruments and clear path for pollutants to inlet.​
Radiation shield: ​
To keep sun from beating on thermometer and artificially raising temperature.​
​

## Slide 85: Site Information
Probe Information
  Probe Material
  Probe Dimensions
  Manifold Description
  Manifold Dimensions
  Tubing Material
  Tubing Dimensions
  Residence Time
    Probe, Manifold, Tubing, Total
Notes: Probe material description and dimensions; and calculate the residence time in the probe.​
The inlet is often a glass cane, sometimes the end of the probe tubing is curled down to act as the inlet.​
The probe tubing is most often Teflon tubing, occasionally glass. For hydrocarbons glass should be used Teflon can adsorb/desorb hydrocarbons.​
Dimensions can vary, but too large diameter or too long length will result in excessive residence time.​
Residence time must be less than ten seconds from the inlet to the distribution manifold and less than 20 seconds to the instruments. ​

## Slide 86: LUNCH BREAK
Returning at 1pm

## Slide 87
Gas Inlet

## Slide 88
Instrument Manifold

## Slide 89
Instrument Manifold

## Slide 90: Non-Dispersive IR Analyzer
Notes: Used mainly for CO2. Also used for SO2, NOx & CO​
​
 IR light generated by a lamp​
 Filtered to provide a  in which the pollutant of concern absorbs the light​
​
 The IR light passes through two separate cells​
 One is filled with a gas which does not absorb IR light (nitrogen or argon) ​
 Other has a continuously flowing stream of sample gas​
​
 Due to the light absorption in the sample cell​
 Less light reaches the detector​
 Difference in energy levels is sensed by detector   Converted to an electrical signal ​

## Slide 91
NOx & CO Analyzer

## Slide 92
Let’s Discuss GFC CO Analyzer
Notes: Gas Filter Correlation Analyzer​
Gas filter coorelation—CO—splits light through 0% CO then 100% then thru the sample​
​

## Slide 93: Gas Filter Correlation Analyzer
Notes: Gas Filter Correlation Analyzer​
​
If desired :​
​
 By inserting gas filters​
 Corresponding to the gases to be measured​
​
 Upto 8 different gasses can be measured by a single instrument​
​
 SO2, NOx, CO2, CO, NH3, H2O & HCl​
(T.Q. #14)

## Slide 94
Looking Inside a GFC CO Analyzer

## Slide 95: Chemiluminescence NOx Analyzer
Notes: Chemiluminescent Analyzers ​
​
​
Chemiluminescence (sometimes "chemoluminescence") is the emission of light (luminescence), as the result of a chemical reaction. There may also be limited emission of heat. ​
◊ In theory, one photon of light should be given off for each molecule of reactant,​
​
 Reaction between nitric oxide (NO) & O3 ​
 Which generate IR (500 - 3000 nm) light ​
​
 Gas sample containing NO is ​
 Bought into a reaction chamber ​
​
O3 generated within the instrument is quantitatively mixed with the sample gas in order to initiate the chemiluminescent reaction. ​
​
The sample gas flow rate is carefully controlled  Electrical signal is  to the NO Conc.​
(T.Q. #13)

## Slide 96
Thermo Environmental Chemiluminescence NOx Analyzer
Notes: A very common NOx analyzer​

## Slide 97
Typical NOx Analyzer

## Slide 98
Chemiluminescence NOx Analyzer

## Slide 99: Non-Dispersive UV Analyzer
Notes: Light from a UV source passes thru the sample cell ​
 to a set of bandpass filters rotating on a wheel ​
 Then to the light detector ​
​
 Measurement filter allows light centered in a narrow band at 285 nm to pass thru the detector ​
 Ref. Filter allows light in the region of 578 nm to pass thru​
​
 SO2 will absorb light at 285 nm ​
 No absorption occurs at 578 nm​
​
The signals are amplified​
Transmittance gives an output  to SO2 Conc.​
​

## Slide 100: Fluorescence SO2 Analyzer
Notes: Fluorescence ​
Fluorescence spectroscopy aka fluorometry or spectrofluorometry, is a type of electromagnetic spectroscopy which analyzes fluorescence from a sample. It involves using a beam of light, usually ultraviolet light, that excites the electrons in molecules of certain compounds and causes them to emit light; typically, but not necessarily, visible lig​
​
Process in which light energy is :​
 Absorbed at one  and​
​
 Emitted at a different ​
​
Molecule emits light when : ​
@ High energy state  Low energy state ​

## Slide 101
Typical CO & O2 Analyzer

## Slide 102: Calibrations and Zero Air
Pure (zero) air generators
Certified cylinder gases
Dilution calibration systems
Calibration is the process of establishing the relationship between the output of a measurement process and a known input
Notes: Gaseous standards for CO, SO2, and NO2 must be traceable to NIST or CRM.​
O3 test concentrations measured by UV photometer.​
Flow measuring instruments traceable to authoritative volume.​
​
Calibration gases and zero air needed for:​
Daily auto zero/spans​
Bi-weekly precision checks (O3, SO2, NO2, CO)​
Weekly zero/spans​
Full calibrations: done every six months by outside staff member with independent calibration gas and zero air​
Equipment:​
Aadco pure air generator: cleans ambient air, periodically checked vs certified zero air​
Certified cylinder gases: for span checks and calibrations​
Dilution calibration systems: use concentrated cylinder gases and dilute to the desired concentration; useful for audits, but must be calibrated.​

## Slide 103
Calibration Gases

## Slide 104
Calibration Gases
Notes: Gaseous standards for CO, SO2, and NO2 must be traceable to NIST or CRM.​
O3 test concentrations measured by UV photometer.​
Flow measuring instruments traceable to authoritative volume.​
​
Calibration gases and zero air needed for:​
Daily auto zero/spans​
Bi-weekly precision checks (O3, SO2, NO2, CO)​
Weekly zero/spans​
Full calibrations: done every six months by outside staff member with independent calibration gas and zero air​
Equipment:​
Aadco pure air generator: cleans ambient air, periodically checked vs certified zero air​
Certified cylinder gases: for span checks and calibrations​
Dilution calibration systems: use concentrated cylinder gases and dilute to the desired concentration; useful for audits, but must be calibrated.​
​

## Slide 106: Particulate Properties
Collected Mass
Inertial Properties
Particle Size
Optical Density
  Haze and Opacity in the Air
  Density of Collected Deposit
Notes: Particulate properties:​
Mass is the most common property of particulate important to ambient monitoring. ​
Inertial properties: Large or heavy particles tend to go in a straight line; small or light particles follow gas flow (can be used as a measurement tool).​
Particle size: Fine particles penetrate deeply into the lungs leading to health effects; coarse particles penetrate less. ​
Large particles deposit from the air quickly, small particles remain airborne longer.​
Optical density: Particles with a diameter about the same size as the wavelength of visible light will be optically active - visible as a haze - what people see as pollution - people don’t think much about pollution unless they see a haze. Transmissometers measure the ‘opacity’ of the air.​
Density of a deposit on a filter: Faster darkening of the filter can be a measure of the particulate concentration or haze value. The ARB uses this in the coefficient of haze monitors. Also used in B-gauged instruments which use B radiation rather than visible light to measure the deposits

## Slide 107: Measures of Particulate Matter in the Atmosphere
Total Suspended Particles
PM 10
PM 2.5
Aerosols
Notes: Particulate NAAQS have been moving toward finer particulate. The target center is the particulate matter most responsible for health effects.​
PM Sampling Systems - In 1971, the Environmental Protection Agency (EPA) promulgated primary and secondary national ambient air quality standards for particulate matter measured as "total suspended particulate" or "TSP". In accordance with the Clean Air Act, the EPA has reviewed and revised the health and welfare criteria upon which these primary and secondary particulate matter standards were based.​
On March 20, 1984 the EPA proposed changes in the standards based on it's review and revision of the criteria. On July 1, 1987 the EPA promulgated its final decisions regarding these changes. The final decisions include:​
1- Replacing TSP as the indicator for particulate matter for the ambient standards with a new indicator that includes only those particles with an aerodynamic diameter less than or equal to a nominal 10 micrometers (PM 10); and​
2- a new Federal Reference Method for measurement of PM 10 in the ambient air. The Reference Method for the measurement of atmospheric particulate matter is known as PM 10 is based on selection of PM 10 particles by internal separation, followed by filtration and gravimetric determination of the PM 10 mass on the filter substrate. ​
In 1997 the new PM2.5 standard was promulgated.​
​

## Slide 108: Particulate
Total Suspended Particulate (TSP) Samplers
    Lead
PM10 and/or PM2.5 samplers
    Size Selective Inlet
    BAM
    TEOM
Visibility Samplers
    Nephelometer
    Optical Test Tape Sampler
Notes: ​
Total Suspended Particulate (TSP): General Metals (GMW) HiVol most common; TSP not sampled often anymore ​
PM10: Sierra Anderson 1200 most common sampler - HiVol with a Size Selective Inlet (SSI) ​
Dichot - gives PM10 to PM2.5 and PM2.5 and less size fractions, sum of the two size fractions = PM10 - Sierra Anderson 244 most common​
TEOM - Tapered Element Oscillation Microbalance - gives continuous PM10 data - made by Rupprecht and Patashnick model 1400A - The TEOM is a monitor rather than a sampler.​
​
​

## Slide 109
Size Selective Inlet (SSI) Sampler
Notes: The size-selective inlet is an aerodynamic filter which removes particles that are not considered harmful to humans. The suspended particles which pass through the filter are collected on a pre-weighed filter over a pre-determined period. After sampling, the filter is removed and re-weighed to determine the concentration of suspended particulate. Results are expressed as micrograms per cubic meter (µg/m3) of air sampled at standard temperature and pressure conditions. ​
The concept for a Size-selective Inlet originated during the late 1970's when the USEPA began to develop health-based ambient air quality standards for particles. The USEPA concluded that particles capable of penetration through the upper respiratory tract could deposit in the tracheobronchial and alveolar regions of the human body and cause adverse health effects. Researchers in the field identified that the particle size range capable of deposition within the human thoracic region was significant for particles less than 10 um in diameter (PM 10)

## Slide 110: PM10 - Size Selective Inlet (SSI)
Buffer Chamber
Acceleration Nozzle
Impaction Chamber
Vent Tubes
Filter Cassette
Filter
Filter Support Screen
Air Flow
Acceleration Nozzle
Impaction Chamber
Motor Inlet
Notes: ​
PM10 most commonly measured with High Volume sampler (HiVol) fitted with a Size Selective Inlet (SSI)​
​
(Additional information: ARB QA Manual Vol II, appendix P)​
​
Inertial separation occurs in acceleration tubes - PM10 and smaller passes to filter; larger particles retained in sampler head. ​
​
Around 40 cfm is pulled through the tubes. At that rate they go critical (or speed of sound), thus the flow is maintained regardless of filter loading, ambient temperature, or pressure conditions. ​
​
(When air reaches the speed of sound a sound pressure barrier is setup that will not allow more air through under existing temp and pressure. See Graseby SSI handout for more info).​

## Slide 111
Inside SSI Head
Notes: Note acceleration nozzles.​
​
The volume of air and its velocity is controlled by the mass flow controller or by a volumetric flow control orifice. Standard flow must be 40 SCFM +/- 10%.​

## Slide 112
SSI Filter Area
Notes: Note filter and timers/flow controllers.​
​
The sample is collected for 24 hours, midnight to midnight, on a once every sixth day basis. ​
​
The sample is taken on a pre-inspected and pre-weighed, high-purity quartz, 8"x10" micro-fiber filter. ​
​

## Slide 113
PM2.5 Sampler
Notes: (TQ #10)​
​
The Dichotomous Sampler operates at a flow rate of 1 CMH (16.7 LPM) and is available with built-in mechanical 6-day or 7-day timers, or an electronic digital programmable timer. Separate sampling and control modules allow for flexibility in deployment. The weather resistant control module houses the diaphragm vacuum pump, precision flow meters, flow selector valves and vacuum gauges, as well as the flow controller and flow recorder. ​
​
The Dichotomous (Dichot) Sampler separates particles into two distinct size fractions 2.5 to 10 microns ("coarse" particles) and less than 2.5 microns ("Fine" particles). ​
​
Fine fraction contributes to visible haze and is more important in health effects and related to industrial processes.​
​
Coarse fraction related to agricultural, aggregate processes, and natural processes. ​

## Slide 115: PM10 Size Selective Inlet
Screen
Upper Plenum
Acceleration Jet
Deflector Cone
Flow to Virtual Impactor
Vent Jet
Ambient Air Flow
Middle Plenum Particles > 10 um
Lower Plenum Particles <10 um
Notes: The unique design of the PM 10 inlet allows only particles smaller than 10 microns to enter the virtual impactor, where they are separated into two size fractions and collected on Teflon membrane filters.

## Slide 117: Virtual Impactor
From Inlet
To Control Module
Coarse Particles >2.5 um
Virtual Impactor Nozzle
Fine Filter Cassette
15.03 LPM
Coarse Filter Cassette
1.67 LPM
PM 10 to 2.5 Coarse Particle Filter
PM 2.5 to 0 Fine Particle Filter
Fine Particles <2.5 um
Virtual Impactor Receiver Tube
Inlet Drift Tube
Notes: Commercially available PM10 dichot samplers operate at an actual flow rate of 16.7 LPM (liters per minute) and a coarse flow of 1.67 LPM. Ninety percent of the air (15.03 LPM) flows through the fine particulate filter while the remaining 10% flows through the coarse.​
​
The virtual impactor ensures that the particulate deposit is uniformly collected on two Teflon membrane filters which eliminate artifact formation, making the Dichot ideal for X-ray fluorescence and other elemental analysis techniques of air sampling. The virtual impactor eliminates possible problems of particle bounce and reintrainment often experienced in other particle size impactor methods.

## Slide 118: EPA-WINS PM2.5 Impactor
PM-10 Aerosol from Inlet
Collection Cup with Antispill Ring
Impaction Filter
Nozzle
T
PM-2.5 Aerosol to Filter
S
W
Notes: The EPA has designed and tested a single jet impactor with an oil impregnated substrate known as the Well Impactor Ninety Six (WINS).​
​
Following the normal PM10 inlet is a 12 inch length of sample tube that supports the PM10 inlet above the sample shelter and transitions into the PM2.5 particle separator or impactor. ​
​
This impactor operates at a volumetric sample rate of 16.67 LPM and has a sharp aerodynamic cut-point at 2.5 micrometers. The exit of the single jet impacts the PM10 fraction of ambient particles onto a 37mm glass fiber substrate that is impregnated with 1 ml of diffusion pump oil. The PM 2.5 particles and smaller exit the impactor well and are collected on the 47 mm PTFE filter media.​
​

## Slide 119
Tapered Element Oscillating Microbalance (TEOM) Inlet
Notes: The Series 1400a monitor incorporates an inertial balance patented by R&P. It measures the mass collected on an exchangeable filter cartridge by monitoring the corresponding frequency changes of a tapered element. The sample flow passes through the filter, where particulate matter collects, and then continues through the hollow tapered element on its way to an electronic flow control system and vacuum pump. ​
​
As more mass collects on the exchangeable filter, the tube's natural frequency of oscillation decreases. A direct relationship exists between the tube's change in frequency and mass on the filter. This inertial methodology is a more fundamental means of measuring mass than the use of gravity. The tapered element technology enables highly precise and accurate mass determinations, providing more resolution than conventional microbalances. ​
​
The TEOM mass transducer does not require recalibration because it is specially designed and constructed from non-fatiguing materials. Calibration may be verified, however, using an optional calibration kit that contains a filter of known mass

## Slide 120: PM10 - TEOM
Flow
To Flow Controller
Microprocessor
Electronic Feedback System
Tapered Element
Filter Cartridge
Flow
Sampling Head
Heated Air Inlet
Notes: Detail of oscillating microbalance​
Acts as a tuning fork; as mass on filter increases the vibrational frequency changes, measure the frequency to determine the particulate concentration.​
M = K/f2​
K = constant, f = frequency​
As particles are collected on the filter, the oscillating mass changes and results in a change of the oscillating frequency. An electronic control system maintains the tapered tube in oscillation and continuously measures this oscillating frequency and its changes. To calibrate the system, the restoring force constant (k in equation ) is determined by placing a gravimetrically determined calibration mass on the filter and recording the frequency change due to this mass.​
​

## Slide 121: PM10-TEOM
PM10 Inlet
Auxiliary Flow Controller
Main Flow Controller
Mass Transducer
TEOM Sensor Unit
Sample Tube
Flow Splitter
Inline Filters
TEOM Filter
Bypass Flow Line
Air Tubes
Bypass Fine Particle Filter Assembly
TEOM Control Unit
Vacuum Pump
Notes: Can see diurnal patterns of particulate concentrations, weather frontal passages, etc.​
​
Heated inlet - can loose volatile particulate (light organics, nitrates, etc.) and bound water.​
​
Because the restoring force constant k is a function of temperature, the sampling apparatus (tapered tube, filter) and sampled air are kept at a constant temperature of 50 °C, which is above typical ambient temperatures. Filter lifetimes are usually two to four weeks. After a filter change, the new filter is equilibrated for one to two hours prior to acquiring data.​
​
The 50 °C temperature prevents water vapor condensation and provides a standard sample condition, but it volatilizes most of the ammonium nitrate and some of the volatile organic compounds in atmospheric particles. As a consequence, monitored sites and seasons having high levels of ammonium nitrate and/or organic particulate mass show a poor correlation between time-integrated TEOM and collocated filter measurements (Allen et al., 1997). Meyer et al. (1992) showed ~30% more PM10 mass was obtained with a 30 °C TEOM sampling beside a 50 °C TEOM in the woodburning-dominated environment at Mammoth Lakes, CA.​
​
​

## Slide 123
BAM Sampler
Notes: Uses dichot head to get particle size required (usually PM10 but could be 2.5 or TSP). Particulate deposited on the tape is treated with a low level of beta radiation. An increase in particles collected on the tape causes a lower beta-ray measurement in the measurement chamber.​
​
Ambient air is pulled through the PM-10 sample inlet and dust particles are deposited on the glass fiber filter tape. A low level of beta radiation is emitted from the source which passes through the filter tape and deposited dust. The increase of particles collected on the tape causes a lower beta-ray measurement in the measuring chamber. Them is direct mass collection and simultaneous measurement at this single filter-spot position which results in continuous observation of the increasing dust mass and corresponding concentration. A compensation chamber receives an equal portion of the beta-ray and is used as a reference by comparing the sample measurement. In the measuring chamber with transmitted energy through a compensation chamber foil which exhibits the same absorptivity as a clean filter tape. As particulate matter is collected on the filter, the differential reading changes and the signal is converted by an on-board computer to PM 10 concentrations. Ionization chamber are used as radiation detectors because of their excellent temperature stability and practical unlimited life time. Output from the instrument can be through visual display1 on board computer, telemetry, printout, or storage in an auxiliary disc.

## Slide 124
Meteorological Instruments
Notes: Met instruments - wind speed/direction and relative humidity; wind speed/direction and gas sampling inlet.​
​
The ARB Quality Assurance Section has conducted performance audits of meteorological sensors since 1993.​
​
The parameters audited include wind speed, wind direction, ambient temperature, relative humidity, barometric pressure, and total solar radiation. (see parameters next page).​
​
Note: Not all of the above mentioned instruments are located at each met site. (See ARB chart.​
​
​

## Slide 125: Meteorological Instruments
Wind speed
Wind direction
Atmospheric pressure
Temperature
Relative humidity, dew pt
Solar radiation
Notes: ​
Wind speed and direction: wind speed measured by an anemometer, direction by a wind vane; used for pollutant transport and dispersion​
Vertical wind speed: measured by sensitive anemometer; used for pollutant dispersion (not measured often)​
Temperature: measured by thermometer or thermocouple; used for pollutant dispersion and general meteorological measurements​
Relative humidity/dew point, atmospheric pressure: measured by hygrometer and barometer respectively; used for general meteorological measurements; if visibility is measured, the results are only valid for limited humidity range (10-70%); Relative Humidity or Dew Point needed.​
Rain fall: measured by rain gauge; needed if acid deposition is monitored at the station​
​

## Slide 126
Meteorological Instruments
Notes: Solar Radiation - left Temperature - right​
Parameters Criteria ​
Wind Speed -- Starting Threshold: less than 0.5 m/s​ Accuracy: +/- 0.25 m/s at speeds less than 5.0 m/s +/- 5% above 5.0 m/s ​
Wind Direction -- Starting Threshold: less than 0.5 m/s​ Accuracy: +/- 5 degrees ​
Ambient Temperature -- Accuracy: +/- 0.5 degrees Celsius​
Relative Humidity -- Accuracy: +/- 1.5 degrees Celsius ​
Barometric Pressure -- Accuracy +/- 10.0 Millibars ​
Total Solar Radiation -- Accuracy: +/- 5 %​

## Slide 127
3D Wind Sensor
Notes: On left acid deposition instruments - dry; buckets for wet acid deposition.​
​
The wet acid deposition monitoring network quality assurance program is designed to provide an assessment of data quality being generated by the network. The Wet Acid Deposition Program data quality objectives are met by conducting performance audits, ensuring that the siting criteria are met, and conducting bucket rinses. ​
​
Performance audits are conducted of the analyzing laboratory and each field network site, thus assuring the accuracy of the data generated. The ARB Quality Assurance Section has conducted performance audits of Acid Deposition since 1986. ​
​

## Slide 128
DATA

## Slide 129: Data Handling
Data loggers
  Strip charts
  Computers
  Temporary data storage
  On-line data retrieval
Notes: Data logging is recording and storage of data.​
​
Strip charts: simple form a data logging.​
Computers: data storage and manipulation.​
Data loggers: temporary data storage and interface between the monitors and data storage. Data loggers allow for basic data manipulation and on-line data retrieval.​
​
​
​
​

## Slide 130
Data Acquisition System
Notes:  Used for data conversion​
 Auto controlling, Alarming​
 Recording, Editing, ​
 Calculating and Reporting​
​
All measured data  are not recorded​
​
All recorded data  are not reported​
​
Note : Part 75 sources report ALL data  Quarterly​

## Slide 131: Site Survey Data
Quality Assurance Procedures and Plans
Cleaning Schedule
Calibrations
Station Temperature Control
In-Line Filters
Notes: Miscellaneous items​
Cleaned, serviced, and calibrated on a schedule.​
Stations must have a written QA plan.​
​
Instruments must have a particulate filter: changed on the schedule specified by the method (Additional information: ARB QA Manual Vol II).​
​
Station must be temperature controlled: instruments must operate between 20oC and 30oC.​
​
​

## Slide 132: Documentation
Instrument Log
  Stays with Instrument
  Documents Acceptance Tests
  Documents Routine Maintenance
  Documents Repairs
  Documents Calibrations
  Other Instrument Specific Information
    i.e. Location, History, etc.
Notes: Instrument log documents everything dealing with a specific piece of equipment.​
Calibrations​
Maintenance​
Repairs, etc.​
Must be up to date.​

## Slide 133: Documentation
Station Log
  Stays at Station
  Documents Conditions that may Influence Data
    Nearby Construction
    Changes in Traffic Patterns and Flow
  Documents Alterations of Sampling Train
    Probe and Equipment Changes
  Contains Completed Site Reports
Notes: Station log documents activities at the station and in the area that can influence the data values and quality.​

## Slide 134
Typical Monitoring Station
Notes: A monitoring site could be a full one like this one and the previous slide or be very limited like the next one.​
​
Met, PM10, Teom, all gases​

## Slide 135
Quality Assurance

## Slide 136: Quality Assurance
Valid Data
Air Monitoring Network
Data
Quality Assurance Program
Notes: QA is the key for a successful AAM program.​
​
This is One Section that is Near and Dear to my Heart!!!​

## Slide 137: Quality Assurance
Field QA
  Daily and Weekly Zero and Span Checks
  Semi-Annual Multipoint Calibrations
  External Audits
    Agency Audits
    EPA NPAP (National Pollutant
    Audit Program)
Notes: Routine zero and span checks, calibrations, audits to assess instrument operation. (Additional information: ARB QA Manual Vol II)​
In general:​
Daily sample flow checks​
Weekly zero and span (usually done daily)​
Semiannual multipoint calibration (usually done by ARB).​
​
Each station must be audited annually: ARB QA section does this at most California sites.​
EPA National Pollutant Audit Program (NPAP): EPA audits a sample of the stations each quarter.​

## Slide 138: Data Handling
Data review and editing
  Complete data set
  Reviewed for accuracy
  Reviewed for consistency
Notes: After retrieval data are reviewed to determine validity:​
​
Must pass zero/spans, calibrations, and audits.​
​
Data editing: Data may be edited to apply calibrations, corrections, etc. after retrieval.​
​
Data must be complete:​
​
At least 30 minutes to make hourly average (if greater than 25% of scale - 45 minutes).​
​
Each of three eight hour periods (0000-0759, 0800-1559, 1600-2359) must contain at least six valid hours.​
​
Data must be logical for the station: If the data are not consistent day to day and year to year, the reviewer should find out what is different.​
​
​

## Slide 139: Data Handling
Data Processing
  Upload to AQS (formerly AIRS)
  Air Quality Data Actions
    Data Deletion
    Data Correction
    Links Data to Field QA
AIRS
Notes: Validated data are incorporated into ‘data for record’ database or delivered to end user. ​
Data for record is uploaded to the EPA’s Aerometric Information Retrieval System (AIRS).​
​
Air Quality Data Action (AQDAs) requests link field QA to the data - instruments that fail an audit or are in the warning range generate a request to repair/recalibrate the instrument and correct or delete the data (some failed data can be corrected depending on what the error was, other must be deleted).​

## Slide 140: Station Inspection
Review Siting
Examine Instruments
  Condition, Zero/Spans, Calibration, Audit Results
Examine Gases
  Certification
Review Logs
Evaluate Overall Station Cleanliness and Operation
Notes: (TQ #24)​
​
When inspecting a monitoring station bring together all that has been covered in class.​
​
Review siting: Does the station meet the siting requirements for the purpose and scale?​
​
Examine the instruments: Are they in good condition? Daily or weekly zero/spans done? Zero/spans at proper levels? Do the instruments drift? When was the last calibration? When was the last audit? Did it pass the audit? Were any problems pointed out by the audit corrected?​
​
Examine span gases: Are the gases certified?​
​
Review the instrument and station logs: Up to date? Any unusual activities?​
​
Evaluate the overall station: Appear well operated? Station neatly maintained? (if a station is neat and clean it is more likely that the operator cares)​

## Slide 141: California ARB Audit Van

## Slide 142: ARB Audit Van Instrumentation
Notes: The basic process for an Air Resources Board audit is to use a gas calibrator to generate three audit concentrations. The exact dilution ratio is determined by analyzing the output of the calibrator with a carbon monoxide analyzer. Carbon monoxide is an unreactive tracer gas which is present in compressed gas cylinders used by the QAS to perform audits. The dilution ratio is calculated by dividing the carbon monoxide analyzer's response by the known concentration of carbon monoxide in a gas cylinder traceable to the National Institute of Standards and Technology (NIST).​
The QAS uses six gas cylinders to perform audits for carbon monoxide, total hydrocarbons, methane, non-methane hydrocarbons, sulfur dioxide, nitrogen dioxide, and hydrogen sulfide. An array of solenoids allow gas to flow from the cylinders to the gas calibrator. The solenoids are actuated by a solid state relay card which is controlled by the Automated Audit Program using assembly language routines.​
The Quality Assurance Section (QAS) began the development of an automated audit with a feasibility study in 1992. The study identified a new generation of computerized ambient calibration systems which can be controlled remotely through a RS-232 serial interface.​
To obtain data from the audit van's instruments, instrument output voltages are applied to a 12 bit analog to digital converter (A/D). The output of the A/D is then latched and converted into a pollution concentration which is brought into the Automated Audit Program by another assembly language routine.​
Ambient data for record for air monitoring stations is continually recorded by an instrument called a data logger. The data logger also utilizes an A/D to convert instrument voltages into pollutant concentrations. Data loggers can transmit instantaneous and recorded pollutant concentrations through RS-232 ports. The QAS developed assembly language routines to poll data loggers through their RS-232 ports. Currently, a 150 foot ultra low capacitance shielded RS-232 cable connects the data logger to the computer in the audit van. ​
​

## Slide 143
Accuracy
&
Precision
Notes: Source test samples should be both accurate and precise. ​
​
If precise, but not accurate - can include a ‘calibration factor’ to remove the bias. ​
​
If not precise - may mean you need to work on your methods and procedures.​
​

## Slide 144: Safety
Compressed Gas Cylinders
Hazardous Gases
Electrical Hazards
Heights
Notes: ​
Safety should always be considered​
​
Key safety considerations:​
compressed gas cylinders​
hazardous gases (H2S, CO, etc.)​
heights (to probe inlet and met tower)​
​

## Slide 145
Notes: AIR NOW is a graphic display of the Air Quality Index both today and forecast for tomorrow this is based on the weather and monitoring trends ​

## Slide 146: www.epa.gov/outdoor-air-quality-data

## Slide 147: www.epa.gov/outdoor-air-quality-data

## Slide 151: The future
Greenhouse Gases
Real time Particulate Speciation
Satellite Stations
???????

## Slide 152: The Web
https://www.epa.gov/outdoor-air-quality-data Monitoring data
http://www.airnow.gov
  AQI
https://www.epa.gov/technical-air-pollution-resources
  NAAQS
  Air monitoring regulations and information
https://www.epa.gov/green-book
  Non attainment Areas

## Slide 154
Questions???